# 20/03/2013 - Stage Calibration

Outline

SI Measurements

Analysis

Coarse (NP) delay stage

Fine (PI) delay stage

Wedge (Ag) dispersion

Wedge calibration 22/03/2013

# Outline

- Calibration of the coarse delay introduced by the Newport stage, fine delay introduced by the PI stage and phase introduced by the Agilis stage used to insert the wedge.

- The Newport (NP) stage and PI stage are placed in the probe pulse (PP) arm of the interferometer.

- NP = 0 corresponds to PP before TP

- PI arranged with overhang towards PP after TP.

- The Agilis (Ag) stage is placed in the test pulse (TP) arm of the interferometer.

- Ag = 0 corresponds to maximum wedge insertion.

- The sign of the phase difference not only depends on which sideband is selected, but also the sign off the delay between the two pulses (i.e. which pulse arrives first) , which are equivalent things

  - For example I(w) = I(w)cos(wt + dp) --> I(w)cos(-wt + dp).

  - Therefore the phase may need to be conjugated to match the physical situation as described above.

  - Since only one stage is translated between measurements, it is possible to track which pulse arrives first.

# SI Measurements

- Settings remain unchanged unless specifically stated

| Measurement | NP (mm) | PI (um) | Ag (mm) | Notes |
| --- | --- | --- | --- | --- |
| 01 | 6.065 | 0 | ??? | Settings same as previous scan. |
| 02 | 6.065 | 19 | 13.5 | Agilis stage homed, range = 0-27mm |
| 03 |  | 0 |  |  |
| 04 |  | 38 |  |  |
| 05 | 6.035 |  |  |  |
| 06 | 6.075 |  |  | Stage uses backlash compensation for +ve movement. |
| 07 | 6.065 |  |  |  |
| 08 |  |  | 13 |  |
| 09 |  |  | 12 |  |
| 10 |  |  | 15 | Analysis below predicts position to be 14mm |
| 11 |  |  | 8 |  |

## Analysis

- In the Takeda algorithm, I take the positive sideband:

  - D = ifft(fft(I).*ifftshift(gauss1D((1:Nw)', n0, dn, 4), 1));

  - n0>N/2

### Coarse (NP) delay stage

- Measurements 04-07

- As delay increases, PP moves from leading to trailing edge of TP, therefore delay corresponds to physical fine of TP.

- Larger delay at larger NP.

- Sign of phase slope is same as sign of delay.

- Relative accuracy of NP stage ~200nm/1.7fs (equal to step size).

- Define positive delay as moving towards trailing edge of TP, which corresponds to smaller NP positions.

Calibration of coarse delay stage ¦ Measured interferometric intensity (dashed), and measured (solid) and theoretical (dotted) phase difference as a function of coarse stage position. Theoretical curve calculated using stage position set-point and spectrometer spectral axis. Measured data averaged over 1,000 samples. RMS variation of measured phase difference, and difference between measured and theoretical curves is negligible (cannot be resolved on graph over region of significant spectral intensity).

### Fine (PI) delay stage

- Measurements 02-04.

- Direction of travel coincides with coarse delay stage:

  - Delay equals physical time of TP.

  - Positive translation = positive delay = positive phase slope

Calibration of fine delay stage ¦ Measured interferometric intensity (dashed), and measured (solid) and theoretical (dotted) phase difference as a function of fine stage position. Theoretical curve calculated using stage position set-point and spectrometer spectral axis. Measured data averaged over 1,000 samples. RMS variation of measured phase difference, and difference between measured and theoretical curves is negligible (cannot be resolved on graph over region of significant spectral intensity).

### Wedge (Ag) dispersion

- Measurements 07-11

- Not all measurements seem to give a phase difference that is consistent with the stage displacement.

- Inconsistency between length of fused silica required to provide measured group delay and measured dispersion

  - Dispersion requires ~3 times as much length compared to group delay

  - Dispersion measurements are consistent with ARIES measurements

    - As wedge is inserted, pulse propagates through less air

    - Only the difference in group delay between the additional fused silica and free space is measured

    - Δϕ=nFSω-nairωωLc

    - Thus for nFS~1.5 and nair~1 →  n/(n-1)~3

Calibration of wedge insertion ¦ Measured interferometric intensity (dashed), and measured (solid) and theoretical (dotted) phase difference as a function of wedge stage position. Theoretical curve calculated using stage position set-point and spectrometer spectral axis. Measured data averaged over 1,000 samples. RMS variation of measured phase difference, and difference between measured and theoretical curves is negligible (cannot be resolved on graph over region of significant spectral intensity). The phase differences above correspond to the additional phase added to the probe pulse, hence resulting in negative wedge path-length for negative wedge insertion.

- From the figure above, it appears that measurement 10 (zAg = 15 mm) appears to be incorrect

  - Decreasing the wedge insertion by a factor of 3 should be consistent with the rest of the data

  - Δz → 0.5 mm

  - zAg → 14 mm

- The wedge angle is determined to be ~2.831°.

- The additional amount of fused silica is given as

  - ΔLFS = (-49.4502 ± 0.6799) × 10-3 ΔzAg

- Larger Ag stage positions results in less wedge insertion, increased delay (i.e. PP arriving later wrt. TP), less TP dispersion (i.e. more relative dispersion on the TP).

- The additional TP phase is given by the negative phase difference

Corrected calibration of wedge insertion ¦ Same as previous figure, except the displacement of the penultimate measurement has been reduced by a factor of 3.

### Wedge calibration 22/03/2013

- Measurements taken during this day matches with the calibration done previously.

Wedge calibration ¦ Additional phase introduced by wedge insertion. The group delay is negative compared to calibration above because the PP must have been set to arrive after the TP.