## CoverSheet (hidden)
Sample Type | Sample Name | Thickness (cm) | Resistivity (Ω-cm) |  |  |  | Wafer Lifetime Measurement Software
Default | Test | 0.0525 | 50 |  |  |  | Version | 4.4.8 | © Sinton Instruments 2013
Results of Measurement | [5 empty] | Error Messages
Lifetime at Spec. MCD (µs) | Sheet Resistance (Ω/sq) | Measured Resistivity (Ω-cm) | J0 (A/cm2) | Fit Intercept (µs) | 1 sun Implied Voc (V) | Input values must be updated, Click the Update button above
363.872 | 896.4372 | 62.7506 | 1.1968e-12 | 874.4002 | 0.59746
Note on A4: The Sample Type seelcted from pre-set values on the InputRanges sheet.
Note on B4: The Sample Name is also used for the filename when saving data so avoid special characters: / ? * etc
Note on C4: Thickness of the wafer in cm e.g. A 100 µm wafer is 0.01 cm thick.
Note on D4: Wafer bulk resistivity. An accurate resistivity (not nominal) is required for accurate calculation of Implied Voc.
Note on A7: Lifetime of minority carriers at the specified MCD. (Auger Corrected)
Note on B7: Sheet resistivity of the entire wafer thickness. If wafer is diffused, parallel combination of bulk and emitter diffusion.
Note on C7: Measured wafer resistivity based on Ohm/sq and the width. Only valid for wafers without diffusions.
Note on D7: J0 as extracted from the slope of inverse lifetime vs MCD fit. The extracted J0 is related to the recombination in the heavily diffused regions (see user manual for more details)
Note on E7: Fit intercept used to estimate bulk lifetime from the inverse lifetime vs MCD fit. Only valid for wafers with diffusion. It is usually different to the low level injection lifetime.
Note on F7: The Implied Voc at an intensity of 1 sun. This value depends heavily on the input resistivity.
Note on G7: How to Fix Error Messages: 1. Saturated Reference Cell: Decrease peak light intensity by raising flash height or adding additional filters. Data at lower MCD/intensit is still valid. 2. Choose Specified MCD within Data: The Specified MCD is not between the Min and Max MCD. Ensure Sample and Analysis Parameters are correct and select Specified MCD between Min and Max MCD. To reach higher Max MCD, increase illumination intensity. To reach lower Min MCD, decrease illumination intensity or extend data acquisition time from Settings page. 3. Re-Zero Instrument: The "Zero Instrument" routine has not been run in the past 12 hours OR the air voltage is out of the normal range. Remove samples from the instrument stage and press the "Zero Instrument" Button. 4. Choose Bias Light within Data: Select a light bias value within the range of measured illumination intensities. 5. Input Values must be updated: Click the Update button to update the input values for the selected "Sample Type"
Rescale Graph

## InputRanges (hidden)
Table for input data for wafer types
Selection | Default | p-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
| Sample | Wafer Type | Optical Constant | Spec MCD | Analysis Mode | PC Zoom | Ref Zoom | Fit Range | # of Flashes to Average
1 | Default | p-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
2 | Type 2 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
3 | Type 3 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
4 | Type 4 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
5 | Type 5 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
6 | Type 6 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
7 | Type 7 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
8 | Type 8 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
9 | Type 9 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
10 | Type 10 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
11 | Type 11 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
12 | Type 12 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
13 | Type 13 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
14 | Type 14 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1
15 | Type 15 | n-type | 0.7 | 1000000000000000 | Generalized (1/1) | 5 | 0.2 | 30% | 1

## User
Wafer Lifetime Measurement Software
[10 empty] | Version | 4.6.0 | © Sinton Instruments 2014
Sample Parameters |  |  |  | Analysis Parameters |  |  |  | Measurement Details
Sample Name | Wafer Thickness (cm) | Resistivity (Ω-cm) | Sample Type | Optical Constant | Specified MCD (cm-3) | Bias Light (suns) | Analysis Mode | Date | Time | Log to Summary
nSi_AlOx10B_450 | 0.07 | 40 | n-type | 0.7 | 1000000000000000 | 0 | Generalized (1/1) | 2022-07-07 | 16:55 | Logged
Results of Measurement | [11 empty] | Error Messages
Lifetime at Spec. MCD (µs) | Sheet Resistance (Ω/sq) | Measured Resistivity (Ω-cm) | J0 (A/cm2) | Fit Intercept (µs) | Min MCD (cm-3) | Max MCD (cm-3) | Bias point CD (cm-3) | Trap Density (cm-3) | Doping (cm-3) | 1 sun Implied Voc (V) | Implied FF (%)
363.872 | 896.4372 | 62.7506 | 1.1968e-12 | 874.4002 | 1.3953e+14 | 8318371307832016 | 0 | 0 | 1.1092e+14 | 0.59746 | 73.6699
Lifetime = 363.87 µs @ 1.0E+15 cm-3
DC
0.014674
Space for User Comments
Rescale Graph
Rescale Graph
Zero Instrument
Measure Wafer
Log To Summary
Save Graph Data to .txt File
Generate SEMI PV13 Report
Note on A5: The Sample Name is also used for the filename when saving data so avoid special characters: / ? * etc
Note on B5: Thickness of the wafer in cm e.g. A 100 µm wafer is 0.01 cm thick.
Note on C5: Wafer bulk resistivity. An accurate resistivity (not nominal) is required for accurate calculation of Implied Voc.
Note on D5: Select p-type or n-type from the drop-down menu.
Note on E5: Photogeneration fraction compared to 38 mA/cm² default. See User Manual for more details.
Note on F5: Minority Carrier Density (MCD) for calculating the lifetime. For most materials the lifetime is not a single parameter but a function of minority carrier density.
Note on G5: Light bias corrects traps and removes anomalously high lifetimes in QSS mode. Not necessary for all samples or in transient mode. See user manual for more details on how to recognize a sample with trapping. Typically bias light should be set to 3 suns if trapping is apparent in the signal.
Note on H5: Select Analysis Mode from Drop-down Menu. Use QSS mode for samples with lifetime less than 100us. Set flash to 1/1 for QSS mode. Use transient mode for samples with lifetime greater than 200us, with flash set to 1/64. For samples between 100-200us, use Generalized 1/1 or Generalized 1/64, with the corresponding flash setting.
Note on I5: Date of the measurement
Note on J5: Time of the Measurement
Note on K5: Notes if data has been logged to summary file(s). See Settings sheet for options.
Note on A8: Lifetime of minority carriers at the specified MCD. (Auger Corrected)
Note on B8: Sheet resistivity of the entire wafer thickness. If wafer is diffused, parallel combination of bulk and emitter diffusion.
Note on C8: Measured wafer resistivity based on Ohm/sq and the width. Only valid for wafers without diffusions.
Note on D8: J0 as extracted from the slope of inverse lifetime vs MCD fit. The extracted J0 is related to the recombination in the heavily diffused regions (see user manual for more details)
Note on E8: Fit intercept used to estimate bulk lifetime from the inverse lifetime vs MCD fit. Only valid for wafers with diffusion. It is usually different to the low level injection lifetime.
Note on F8: Minimum Apparent Minority Carrier Density (MCD).
Note on G8: Maximum Apparent Minority Carrier Density (MCD).
Note on H8: Bias Point Carrier Density
Note on I8: The density of electron-hole pairs with the equivalent conductance as the trap density.
Note on J8: The wafer doping of the base region as calculated from the resistivity and dopant type.
Note on K8: The Implied Voc at an intensity of 1 sun. This value depends heavily on the input resistivity.
Note on L8: The Implied FF. This output depend heavily on the input parameters.
Note on M8: How to Fix Error Messages: 1. Saturated Reference Cell: Decrease peak light intensity by raising flash height or adding additional filters. Data at lower MCD/intensit is still valid. 2. Choose Specified MCD within Data: The Specified MCD is not between the Min and Max MCD. Ensure Sample and Analysis Parameters are correct and select Specified MCD between Min and Max MCD. To reach higher Max MCD, increase illumination intensity. To reach lower Min MCD, decrease illumination intensity or extend data acquisition time from Settings page. 3. Re-Zero Instrument: The "Zero Instrument" routine has not been run in the past 12 hours OR the air voltage is out of the normal range. Remove samples from the instrument stage and press the "Zero Instrument" Button. 4. Choose Bias Light within Data: Select a light bias value within the range of measured illumination intensities.

## Calc (hidden)
Lifetime Data Acquisition and Analysis | [8 empty] | Spec MCD | Lifetime | µs | Lifetime = 363.87 µs @ 1.0E+15 cm-3 |  |  |  | Sat. Ref Cell |  | Analysis Mode
© Sinton Instruments 2014 |  |  |  | Date | 2022-07-07 |  |  | Text Labels | 1.0E+15 | 363.87 | cm-3 |  |  |  |  | MCD out of Range |  | Update Cover Sheet | Input values must be updated, Click the Update button above
|  |  |  | Time | 16:55 |  |  |  | Max Valid MCD | [6 empty] | Light Bias out of Range
[9 empty] | 0.0E+00 | [6 empty] | Instrument Calibration
Dark Voltage | Sheet rho Ohms/sq | Resistivity Ohm-cm | Doping (cm-3) | PC/Ref | Max suns | Fit Range Min | Fit Range Max | Min Apparent CD | Max Apparent CD | Time of Max Apparent CD | PC at Time of Max Apparent CD | "Carrier density" at bias point | Traps (equiv e-h pairs) | SpecMCD - Traps | 1/Tau Fit Slope | 1/Tau Fit Intercept | Tau @ specMCD (μs) | Tau @ specMCD (s) | 1/Tau (auger corrected) Fit Slope | 1/Tau (auger corrected) Fit Intercept | Tau (auger corrected) @ specMCD (µs) | Tau (auger corrected) @ specMCD (s) | 1/Tau (auger corrected) @ specMCD | Tau fit (µs) (intercept @ -NA) | ni2/tau fit slope | If h-l-i data, then Jo (A/cm2) | Implied Voc @ 1 sun | arb Jsc | Max Power | Vmp | Jmp | Implied FF | MCD at 0.6V Ref Cell
-0.8454 | 896.4372 | 62.7506 | 1.1092e+14 | 300.9812 | 48.8362 | 700000000000000 | 1300000000000000 | 1.3953e+14 | 8318371307832016 | 0.0004499 | 0.15101 | 0 | 0 | 1000000000000000 | 1.4461e-12 | 1302.1061 | 363.872 | 0.00036387 | 1.4428e-12 | 1303.6812 | 364.0983 | 0.0003641 | 2746.5109 | 874.4002 |  | 1.1968e-12 | 0.59746 | 0.035 | 0.015405 | 0.52549 | 0.029316 | 73.6699 | 0
Time in s | Photovoltage | Reference Voltage | PCD with baseline shift. | Ref (suns) | Delta Suns | Generation (pairs/s) | Conductivity increase | Apparent CD | Apparent CD In Range | Delta N | dn/dt | Suns near bias point | CD near bias point | Minority Carrier Density | Tau (sec) | Tau In Range | 1/Tau in Range | 1/Tau Corrected | 1/Tau corrected in Range | 1/Tau Fit (auger corrected) | ni | ni2/tau | Implied Voc | Implied Suns | Power | Jeq | MCD for Ref Cell>0.6V
4.99e-05 | 1.2987 | 0.2308 | 1.2987 | 27.441 | 27.441 | 4.5564e+18 | 0.036312 | 1.8605e+15 |  | 1.8605e+15 |  |  |  | 1.8605e+15 | [5 empty] | 3988.0433
0.0001499 | 3.5962 | 0.34791 | 3.5962 | 41.3648 | 41.3648 | 6.8683e+18 | 0.10219 | 5467332468520812 |  | 5467332468520812 |  |  |  | 5467332468520812 | [5 empty] | 9192.1109
0.0002499 | 4.7733 | 0.41074 | 4.7733 | 48.8362 | 48.8362 | 8.1089e+18 | 0.13676 | 7470796002349974 |  | 7470796002349974 | 1.5729e+19 |  |  | 7470796002349974 | 7.4624e-05 |  |  | 13307.873 |  | 12082.7675
0.0003499 | 5.2317 | 0.40933 | 5.2317 | 48.6676 | 48.6676 | 8.0809e+18 | 0.15037 | 8280128858683653 |  | 8280128858683653 | 6.1142e+18 |  |  | 8280128858683653 | 7.5737e-05 |  |  | 13089.7263 |  | 13250.497
0.0004499 | 5.2532 | 0.3818 | 5.2532 | 45.3951 | 45.3951 | 7.5375e+18 | 0.15101 | 8318371307832016 |  | 8318371307832016 | 5.1825e+17 |  |  | 8318371307832016 | 7.7625e-05 |  |  | 12767.5189 |  | 13305.6744
0.0005499 | 5.1309 | 0.35719 | 5.1309 | 42.4684 | 42.4684 | 7.0516e+18 | 0.14737 | 8100663770522564 |  | 8100663770522564 | -2.0008e+18 |  |  | 8100663770522564 | 7.8848e-05 |  |  | 12573.6537 |  | 12991.5595 |  |  | 0.7071 | 43.3119 | -1.0472 | -1.4809
0.0006499 | 4.9721 | 0.3325 | 4.9721 | 39.5331 | 39.5331 | 6.5642e+18 | 0.14265 | 7819653846977320 |  | 7819653846977320 | -2.7266e+18 |  |  | 7819653846977320 | 8.1032e-05 |  |  | 12239.2615 |  | 12586.11 |  |  | 0.7053 | 40.6826 | -0.97959 | -1.3889
0.0007499 | 4.8067 | 0.30996 | 4.8067 | 36.8536 | 36.8536 | 6.1193e+18 | 0.13775 | 7529106303172681 |  | 7529106303172681 | -2.8642e+18 |  |  | 7529106303172681 | 8.3395e-05 |  |  | 11896.9994 |  | 12166.8994 |  |  | 0.70337 | 38.0611 | -0.91237 | -1.2971
0.0008499 | 4.6414 | 0.28866 | 4.6414 | 34.3211 | 34.3211 | 5.6988e+18 | 0.13286 | 7240838114708719 |  | 7240838114708719 | -2.8389e+18 |  |  | 7240838114708719 | 8.5945e-05 |  |  | 11548.3387 |  | 11750.9775 |  |  | 0.70138 | 35.518 | -0.84736 | -1.2081
0.0009499 | 4.478 | 0.26829 | 4.478 | 31.8992 | 31.8992 | 5.2966e+18 | 0.12803 | 6957994349400993 |  | 6957994349400993 | -2.7653e+18 |  |  | 6957994349400993 | 8.8714e-05 |  |  | 11191.754 |  | 11342.8821 |  |  | 0.69935 | 33.065 | -0.78486 | -1.1223
0.0010499 | 4.3196 | 0.2496 | 4.3196 | 29.6773 | 29.6773 | 4.9277e+18 | 0.12337 | 6685751356152515 |  | 6685751356152515 | -2.6692e+18 |  |  | 6685751356152515 | 9.1505e-05 |  |  | 10854.2131 |  | 10950.0818 |  |  | 0.69731 | 30.8025 | -0.72736 | -1.0431
0.0011499 | 4.1662 | 0.23242 | 4.1662 | 27.6335 | 27.6335 | 4.5883e+18 | 0.11886 | 6423987082665524 |  | 6423987082665524 | -2.5595e+18 |  |  | 6423987082665524 | 9.4322e-05 |  |  | 10533.5193 |  | 10572.4005 |  |  | 0.69528 | 28.7126 | -0.67438 | -0.96994
0.0012499 | 4.0182 | 0.216 | 4.0182 | 25.682 | 25.682 | 4.2643e+18 | 0.11452 | 6173260391428934 |  | 6173260391428934 | -2.4492e+18 |  |  | 6173260391428934 | 9.742e-05 |  |  | 10201.6039 |  | 10210.6446 |  |  | 0.69325 | 26.7145 | -0.62393 | -0.90001
0.0013499 | 3.8764 | 0.20158 | 3.8764 | 23.9668 | 23.9668 | 3.9795e+18 | 0.11037 | 5934470342486079 |  | 5934470342486079 | -2.3393e+18 |  |  | 5934470342486079 | 0.00010026 |  |  | 9915.3577 |  | 9866.1112 |  |  | 0.69124 | 24.953 | -0.57951 | -0.83835
0.0014499 | 3.7398 | 0.18764 | 3.7398 | 22.3102 | 22.3102 | 3.7044e+18 | 0.10638 | 5705928596947935 |  | 5705928596947935 | -2.2347e+18 |  |  | 5705928596947935 | 0.00010345 |  |  | 9612.2565 |  | 9536.3644 |  |  | 0.68924 | 23.2523 | -0.5368 | -0.77883
0.0015499 | 3.6087 | 0.1746 | 3.6087 | 20.7595 | 20.7595 | 3.447e+18 | 0.10256 | 5487981409533569 |  | 5487981409533569 | -2.1377e+18 |  |  | 5487981409533569 | 0.00010681 |  |  | 9312.2718 |  | 9221.9038 |  |  | 0.68726 | 21.6607 | -0.49698 | -0.72312
0.0016499 | 3.4823 | 0.16314 | 3.4823 | 19.3966 | 19.3966 | 3.2207e+18 | 0.098879 | 5279140938784232 |  | 5279140938784232 | -2.0429e+18 |  |  | 5279140938784232 | 0.00010986 |  |  | 9056.0371 |  | 8920.5825 |  |  | 0.68529 | 20.2578 | -0.4619 | -0.67402
0.0017499 | 3.3606 | 0.15196 | 3.3606 | 18.0671 | 18.0671 | 2.9999e+18 | 0.09534 | 5079038322726614 |  | 5079038322726614 | -1.952e+18 |  |  | 5079038322726614 | 0.00011335 |  |  | 8779.278 |  | 8631.8685 |  |  | 0.68332 | 18.89 | -0.42786 | -0.62615
0.0018499 | 3.2443 | 0.14189 | 3.2443 | 16.87 | 16.87 | 2.8011e+18 | 0.091965 | 4888945383497467 |  | 4888945383497467 | -1.8615e+18 |  |  | 4888945383497467 | 0.00011674 |  |  | 8526.1375 |  | 8357.5968 |  |  | 0.68138 | 17.6548 | -0.39719 | -0.58292
0.0019499 | 3.1324 | 0.13259 | 3.1324 | 15.7649 | 15.7649 | 2.6176e+18 | 0.088725 | 4707078034554299 |  | 4707078034554299 | -1.7734e+18 |  |  | 4707078034554299 | 0.00012018 |  |  | 8284.3576 |  | 8095.1932 |  |  | 0.67946 | 16.5125 | -0.3689 | -0.54294
0.0020499 | 3.0256 | 0.12359 | 3.0256 | 14.6945 | 14.6945 | 2.4399e+18 | 0.085636 | 4534386484947477 |  | 4534386484947477 | -1.6931e+18 |  |  | 4534386484947477 | 0.00012406 |  |  | 8026.2452 |  | 7846.0287 |  |  | 0.67756 | 15.4082 | -0.34169 | -0.50429
0.0021499 | 2.9233 | 0.11597 | 2.9233 | 13.7886 | 13.7886 | 2.2895e+18 | 0.08268 | 4369637826078056 |  | 4369637826078056 | -1.618e+18 |  |  | 4369637826078056 | 0.0001273 |  |  | 7823.657 |  | 7608.3244 |  |  | 0.67568 | 14.4707 | -0.31857 | -0.47148
0.0022499 | 2.8243 | 0.10834 | 2.8243 | 12.8811 | 12.8811 | 2.1388e+18 | 0.079826 | 4211130389768981 |  | 4211130389768981 | -1.5486e+18 |  |  | 4211130389768981 | 0.00013118 |  |  | 7593.9117 |  | 7379.6252 |  |  | 0.67381 | 13.5339 | -0.29559 | -0.43869
0.0023499 | 2.7294 | 0.10114 | 2.7294 | 12.0257 | 12.0257 | 1.9968e+18 | 0.07709 | 4059693690977409 |  | 4059693690977409 | -1.4806e+18 |  |  | 4059693690977409 | 0.0001353 |  |  | 7363.8319 |  | 7161.1278 |  |  | 0.67195 | 12.6499 | -0.27399 | -0.40775
0.0024499 | 2.6383 | 0.095122 | 2.6383 | 11.3097 | 11.3097 | 1.8779e+18 | 0.074469 | 3.9151e+15 |  | 3.9151e+15 | -1.414e+18 |  |  | 3.9151e+15 | 0.00013863 |  |  | 7188.0199 |  | 6952.4336 |  |  | 0.67011 | 11.9058 | -0.25578 | -0.3817
0.0025499 | 2.5512 | 0.088833 | 2.5512 | 10.5619 | 10.5619 | 1.7537e+18 | 0.071967 | 3.7774e+15 |  | 3.7774e+15 | -1.3528e+18 |  |  | 3.7774e+15 | 0.00014305 |  |  | 6966.8708 |  | 6753.8138 |  |  | 0.6683 | 11.1322 | -0.237 | -0.35463
0.0026499 | 2.4673 | 0.083323 | 2.4673 | 9.9069 | 9.9069 | 1.645e+18 | 0.069559 | 3645283068023184 |  | 3645283068023184 | -1.2953e+18 |  |  | 3645283068023184 | 0.00014702 |  |  | 6779.8278 |  | 6563.2039 |  |  | 0.6665 | 10.453 | -0.22051 | -0.33085
0.0027499 | 2.3862 | 0.078244 | 2.3862 | 9.3029 | 9.3029 | 1.5447e+18 | 0.067235 | 3.5182e+15 |  | 3.5182e+15 | -1.2406e+18 |  |  | 3.5182e+15 | 0.00015095 |  |  | 6604.3517 |  | 6379.7939 |  |  | 0.6647 | 9.8259 | -0.20533 | -0.30891
0.0028499 | 2.3087 | 0.073085 | 2.3087 | 8.6895 | 8.6895 | 1.4428e+18 | 0.065014 | 3397013458479484 |  | 3397013458479484 | -1.1881e+18 |  |  | 3397013458479484 | 0.00015583 |  |  | 6398.2327 |  | 6204.9931 |  |  | 0.66293 | 9.1904 | -0.19004 | -0.28666
0.0029499 | 2.2343 | 0.069026 | 2.2343 | 8.2069 | 8.2069 | 1.3627e+18 | 0.062887 | 3.2812e+15 |  | 3.2812e+15 | -1.1399e+18 |  |  | 3.2812e+15 | 0.00015923 |  |  | 6262.3803 |  | 6037.9505 |  |  | 0.66117 | 8.6875 | -0.1779 | -0.26906
0.0030499 | 2.1624 | 0.0647 | 2.1624 | 7.6927 | 7.6927 | 1.2773e+18 | 0.060832 | 3169704805721946 |  | 3169704805721946 | -1.0955e+18 |  |  | 3169704805721946 | 0.00016387 |  |  | 6085.7113 |  | 5877.0254 |  |  | 0.65943 | 8.1545 | -0.16513 | -0.25041
0.0031499 | 2.0926 | 0.060496 | 2.0926 | 7.1928 | 7.1928 | 1.1943e+18 | 0.05884 | 3.0619e+15 |  | 3.0619e+15 | -1.0524e+18 |  |  | 3.0619e+15 | 0.00016903 |  |  | 5900.4519 |  | 5721.4252 |  |  | 0.65768 | 7.6365 | -0.15276 | -0.23228
0.0032499 | 2.0258 | 0.057276 | 2.0258 | 6.8099 | 6.8099 | 1.1307e+18 | 0.056935 | 2958950718217559 |  | 2958950718217559 | -1.009e+18 |  |  | 2958950718217559 | 0.00017241 |  |  | 5785.5742 |  | 5572.9432 |  |  | 0.65595 | 7.2352 | -0.14315 | -0.21823
0.0033499 | 1.9616 | 0.053619 | 1.9616 | 6.3751 | 6.3751 | 1.0585e+18 | 0.055106 | 2.8604e+15 |  | 2.8604e+15 | -9.6902e+17 |  |  | 2.8604e+15 | 0.00017776 |  |  | 5611.8404 |  | 5430.7626 |  |  | 0.65425 | 6.7836 | -0.13244 | -0.20243
0.0034499 | 1.8999 | 0.050447 | 1.8999 | 5.998 | 5.998 | 9.9593e+17 | 0.05335 | 2.7659e+15 |  | 2.7659e+15 | -9.302e+17 |  |  | 2.7659e+15 | 0.00018248 |  |  | 5467.4266 |  | 5294.4758 |  |  | 0.65255 | 6.3902 | -0.12311 | -0.18866
0.0035499 | 1.8396 | 0.047659 | 1.8396 | 5.6665 | 5.6665 | 9.4088e+17 | 0.051633 | 2.6739e+15 |  | 2.6739e+15 | -8.9569e+17 |  |  | 2.6739e+15 | 0.0001865 |  |  | 5349.9802 |  | 5161.5923 |  |  | 0.65085 | 6.0441 | -0.1149 | -0.17654
0.0036499 | 1.7826 | 0.044465 | 1.7826 | 5.2868 | 5.2868 | 8.7782e+17 | 0.050014 | 2587131206289060 |  | 2587131206289060 | -8.6072e+17 |  |  | 2587131206289060 | 0.00019305 |  |  | 5168.7871 |  | 5036.4709 |  |  | 0.64919 | 5.6496 | -0.10565 | -0.16274
0.0037499 | 1.7265 | 0.042314 | 1.7265 | 5.031 | 5.031 | 8.3536e+17 | 0.04842 | 2.502e+15 |  | 2.502e+15 | -8.2634e+17 |  |  | 2.502e+15 | 0.00019608 |  |  | 5089.6146 |  | 4913.6043 |  |  | 0.6475 | 5.3794 | -0.099248 | -0.15328
0.0038499 | 1.6733 | 0.0397 | 1.6733 | 4.7202 | 4.7202 | 7.8376e+17 | 0.046912 | 2421528536526037 |  | 2421528536526037 | -7.9845e+17 |  |  | 2421528536526037 | 0.00020188 |  |  | 4943.76 |  | 4797.5345 |  |  | 0.64586 | 5.0569 | -0.091706 | -0.14199
0.0039499 | 1.6216 | 0.037158 | 1.6216 | 4.4179 | 4.4179 | 7.3357e+17 | 0.045446 | 2.3435e+15 |  | 2.3435e+15 | -7.6872e+17 |  |  | 2.3435e+15 | 0.00020834 |  |  | 4790.685 |  | 4684.9256 |  |  | 0.64421 | 4.742 | -0.084373 | -0.13097
0.0040499 | 1.571 | 0.03537 | 1.571 | 4.2053 | 4.2053 | 6.9827e+17 | 0.04401 | 2267154283863490 |  | 2267154283863490 | -7.4356e+17 |  |  | 2267154283863490 | 0.00021151 |  |  | 4719.3277 |  | 4574.7987 |  |  | 0.64255 | 4.5188 | -0.079136 | -0.12316
0.0041499 | 1.5228 | 0.03314 | 1.5228 | 3.9402 | 3.9402 | 6.5424e+17 | 0.042646 | 2194801256365321 |  | 2194801256365321 | -7.1753e+17 |  |  | 2194801256365321 | 0.00021809 |  |  | 4577.3234 |  | 4470.4056 |  |  | 0.64093 | 4.2427 | -0.072742 | -0.11349
0.0042499 | 1.4756 | 0.03144 | 1.4756 | 3.7381 | 3.7381 | 6.2068e+17 | 0.04131 | 2.1241e+15 |  | 2.1241e+15 | -6.8938e+17 |  |  | 2.1241e+15 | 0.00022227 |  |  | 4491.4815 |  | 4368.3813 |  |  | 0.63928 | 4.0287 | -0.067767 | -0.106
0.0043499 | 1.4302 | 0.029848 | 1.4302 | 3.5488 | 3.5488 | 5.8925e+17 | 0.040027 | 2.0562e+15 |  | 2.0562e+15 | -6.6478e+17 |  |  | 2.0562e+15 | 0.00022639 |  |  | 4410.093 |  | 4270.4989 |  |  | 0.63766 | 3.8291 | -0.063139 | -0.099017
0.0044499 | 1.387 | 0.027914 | 1.387 | 3.3189 | 3.3189 | 5.5108e+17 | 0.038805 | 1.9917e+15 |  | 1.9917e+15 | -6.3927e+17 |  |  | 1.9917e+15 | 0.000234 |  |  | 4266.99 |  | 4177.4198 |  |  | 0.63606 | 3.5884 | -0.057624 | -0.090594
0.0045499 | 1.3445 | 0.02674 | 1.3445 | 3.1793 | 3.1793 | 5.279e+17 | 0.037606 | 1.9286e+15 |  | 1.9286e+15 | -6.1798e+17 |  |  | 1.9286e+15 | 0.00023636 |  |  | 4224.5739 |  | 4086.3038 |  |  | 0.63445 | 3.4398 | -0.054178 | -0.085394
0.0046499 | 1.3039 | 0.025215 | 1.3039 | 2.9979 | 2.9979 | 4.9778e+17 | 0.03646 | 1868287424157004 |  | 1868287424157004 | -5.9814e+17 |  |  | 1868287424157004 | 0.00024234 |  |  | 4120.6201 |  | 3999.3018 |  |  | 0.63286 | 3.2501 | -0.04984 | -0.078753
0.0047499 | 1.2639 | 0.023425 | 1.2639 | 2.7851 | 2.7851 | 4.6245e+17 | 0.035331 | 1.809e+15 |  | 1.809e+15 | -5.7752e+17 |  |  | 1.809e+15 | 0.00025181 |  |  | 3965.8443 |  | 3913.7404 |  |  | 0.63125 | 3.0286 | -0.04482 | -0.071002
0.0048499 | 1.2257 | 0.022517 | 1.2257 | 2.6772 | 2.6772 | 4.4452e+17 | 0.034254 | 1.7525e+15 |  | 1.7525e+15 | -5.5714e+17 |  |  | 1.7525e+15 | 0.00025371 |  |  | 3936.4751 |  | 3832.1919 |  |  | 0.62967 | 2.912 | -0.042139 | -0.066922
0.0049499 | 1.1886 | 0.021107 | 1.1886 | 2.5096 | 2.5096 | 4.1669e+17 | 0.033209 | 1.6977e+15 |  | 1.6977e+15 | -5.3888e+17 |  |  | 1.6977e+15 | 0.00026153 |  |  | 3818.9123 |  | 3753.2294 |  |  | 0.62809 | 2.7367 | -0.038179 | -0.060786
0.0050499 | 1.1531 | 0.019729 | 1.1531 | 2.3457 | 2.3457 | 3.8949e+17 | 0.032207 | 1645338890713628 |  | 1645338890713628 | -5.221e+17 |  |  | 1645338890713628 | 0.00027034 |  |  | 3694.5918 |  | 3677.625 |  |  | 0.62653 | 2.5658 | -0.034336 | -0.054804
0.0051499 | 1.1176 | 0.018978 | 1.1176 | 2.2564 | 2.2564 | 3.7466e+17 | 0.031208 | 1593108604489832 |  | 1593108604489832 | -5.0413e+17 |  |  | 1593108604489832 | 0.00027203 |  |  | 3671.9185 |  | 3602.2656 |  |  | 0.62493 | 2.469 | -0.03213 | -0.051414
0.0052499 | 1.0839 | 0.017658 | 1.0839 | 2.0995 | 2.0995 | 3.4861e+17 | 0.030261 | 1.5437e+15 |  | 1.5437e+15 | -4.8745e+17 |  |  | 1.5437e+15 | 0.00028234 |  |  | 3537.833 |  | 3531.0239 |  |  | 0.62336 | 2.305 | -0.028472 | -0.045675
0.0053499 | 1.0517 | 0.017037 | 1.0517 | 2.0256 | 2.0256 | 3.3634e+17 | 0.029355 | 1.4965e+15 |  | 1.4965e+15 | -4.705e+17 |  |  | 1.4965e+15 | 0.00028368 |  |  | 3521.4402 |  | 3462.8437 |  |  | 0.62182 | 2.224 | -0.026638 | -0.042839
0.0054499 | 1.0199 | 0.016144 | 1.0199 | 1.9194 | 1.9194 | 3.1871e+17 | 0.02846 | 1.4499e+15 |  | 1.4499e+15 | -4.5889e+17 |  |  | 1.4499e+15 | 0.0002893 |  |  | 3453.1247 |  | 3395.6853 |  |  | 0.62025 | 2.1129 | -0.02416 | -0.038951
0.0055499 | 0.98897 | 0.015 | 0.98897 | 1.7834 | 1.7834 | 2.9612e+17 | 0.027591 | 1.4048e+15 |  | 1.4048e+15 | -4.4606e+17 |  |  | 1.4048e+15 | 0.0003004 |  |  | 3325.6657 |  | 3330.5088 |  |  | 0.61868 | 1.9715 | -0.021036 | -0.034001
0.0056499 | 0.95838 | 0.01463 | 0.95838 | 1.7395 | 1.7395 | 2.8883e+17 | 0.026732 | 1.3601e+15 |  | 1.3601e+15 | -4.3355e+17 |  |  | 1.3601e+15 | 0.00029829 |  |  | 3349.3211 |  | 3266.1447 |  |  | 0.61709 | 1.9223 | -0.01992 | -0.03228
0.0057499 | 0.92967 | 0.0137 | 0.92967 | 1.6289 | 1.6289 | 2.7047e+17 | 0.025926 | 1.3183e+15 |  | 1.3183e+15 | -4.205e+17 |  |  | 1.3183e+15 | 0.00030771 |  |  | 3246.9353 |  | 3205.8206 |  |  | 0.61554 | 1.8062 | -0.017369 | -0.028217
0.0058499 | 0.9008 | 0.012726 | 0.9008 | 1.5131 | 1.5131 | 2.5124e+17 | 0.025116 | 1.2764e+15 | 1.2764e+15 | 1.2764e+15 | -4.0564e+17 |  |  | 1.2764e+15 | 0.00031951 | 0.00031951 | 3129.8056 | 3127.1013 | 3127.1013 | 3145.2611 |  |  | 0.61395 | 1.6841 | -0.0147 | -0.023944
0.0059499 | 0.87327 | 0.01241 | 0.87327 | 1.4755 | 1.4755 | 2.4499e+17 | 0.024343 | 1.2364e+15 | 1.2364e+15 | 1.2364e+15 | -3.9531e+17 |  |  | 1.2364e+15 | 0.00031741 | 0.00031741 | 3150.4546 | 3147.9169 | 3147.9169 | 3087.5927 |  |  | 0.61238 | 1.6421 | -0.013763 | -0.022475
0.0060499 | 0.847 | 0.011465 | 0.847 | 1.3632 | 1.3632 | 2.2634e+17 | 0.023607 | 1198299764662150 | 1198299764662150 | 1198299764662150 | -3.8128e+17 |  |  | 1198299764662150 | 0.00033151 | 0.00033151 | 3016.5427 | 3014.159 | 3014.159 | 3032.6237 |  |  | 0.61084 | 1.5239 | -0.0112 | -0.018336
0.0061499 | 0.82034 | 0.010958 | 0.82034 | 1.3029 | 1.3029 | 2.1633e+17 | 0.022859 | 1159718392552213 | 1159718392552213 | 1159718392552213 | -3.6928e+17 |  |  | 1159718392552213 | 0.0003352 | 0.0003352 | 2983.2782 | 2981.0456 | 2981.0456 | 2976.9573 |  |  | 0.60923 | 1.4586 | -0.009778 | -0.01605
0.0062499 | 0.79569 | 0.01047 | 0.79569 | 1.2448 | 1.2448 | 2.0669e+17 | 0.022168 | 1.1241e+15 | 1.1241e+15 | 1.1241e+15 | -3.5872e+17 |  |  | 1.1241e+15 | 0.00033944 | 0.00033944 | 2945.9855 | 2943.8881 | 2943.8881 | 2925.5205 |  |  | 0.6077 | 1.3961 | -0.0084238 | -0.013862
0.0063499 | 0.77131 | 0.0095895 | 0.77131 | 1.1402 | 1.1402 | 1.8931e+17 | 0.021486 | 1.0889e+15 | 1.0889e+15 | 1.0889e+15 | -3.4688e+17 |  |  | 1.0889e+15 | 0.00035685 | 0.00035685 | 2802.3157 | 2800.3475 | 2800.3475 | 2874.7431 |  |  | 0.60614 | 1.2864 | -0.0060758 | -0.010024
0.0064499 | 0.74738 | 0.009548 | 0.74738 | 1.1352 | 1.1352 | 1.885e+17 | 0.020815 | 1.0544e+15 | 1.0544e+15 | 1.0544e+15 | -3.3733e+17 |  |  | 1.0544e+15 | 0.00034796 | 0.00034796 | 2873.8904 | 2872.045 | 2872.045 | 2824.9471 |  |  | 0.60456 | 1.2774 | -0.0058704 | -0.0097103
0.0065499 | 0.7243 | 0.0088995 | 0.7243 | 1.0581 | 1.0581 | 1.7569e+17 | 0.020169 | 1.0211e+15 | 1.0211e+15 | 1.0211e+15 | -3.2698e+17 |  |  | 1.0211e+15 | 0.00035995 | 0.00035995 | 2778.1809 | 2776.45 | 2776.45 | 2776.9993 |  |  | 0.60299 | 1.196 | -0.0041359 | -0.006859
0.0066499 | 0.70215 | 0.0082081 | 0.70215 | 0.97592 | 0.97592 | 1.6204e+17 | 0.019549 | 9.8927e+14 | 9.8927e+14 | 9.8927e+14 | -3.1643e+17 |  |  | 9.8927e+14 | 0.00037596 | 0.00037596 | 2659.8762 | 2658.2516 | 2658.2516 | 2731.034 |  |  | 0.60145 | 1.1093 | -0.0023012 | -0.0038262
0.0067499 | 0.68034 | 0.0082701 | 0.68034 | 0.98329 | 0.98329 | 1.6327e+17 | 0.018939 | 9.5793e+14 | 9.5793e+14 | 9.5793e+14 | -3.0683e+17 |  |  | 9.5793e+14 | 0.00036296 | 0.00036296 | 2755.1346 | 2753.6113 | 2753.6113 | 2685.814 |  |  | 0.59988 | 1.1126 | -0.002365 | -0.0039426
0.0068499 | 0.6593 | 0.0076744 | 0.6593 | 0.91246 | 0.91246 | 1.5151e+17 | 0.018351 | 9.2775e+14 | 9.2775e+14 | 9.2775e+14 | -2.9764e+17 |  |  | 9.2775e+14 | 0.00037682 | 0.00037682 | 2653.7756 | 2652.3468 | 2652.3468 | 2642.2638 |  |  | 0.59832 | 1.0379 | -0.00079455 | -0.001328
0.0069499 | 0.63888 | 0.0071275 | 0.63888 | 0.84744 | 0.84744 | 1.4071e+17 | 0.01778 | 8.9848e+14 | 8.9848e+14 | 8.9848e+14 | -2.8857e+17 |  |  | 8.9848e+14 | 0.00039086 | 0.00039086 | 2558.4747 | 2557.1347 | 2557.1347 | 2600.0349 |  |  | 0.59676 | 0.9691 | 0.0006454 | 0.0010815
0.0070499 | 0.61911 | 0.0070226 | 0.61911 | 0.83496 | 0.83496 | 1.3864e+17 | 0.017227 | 8.7018e+14 | 8.7018e+14 | 8.7018e+14 | -2.8085e+17 |  |  | 8.7018e+14 | 0.0003848 | 0.0003848 | 2598.779 | 2597.522 | 2597.522 | 2559.2036 |  |  | 0.59521 | 0.95336 | 0.00097158 | 0.0016323
0.0071499 | 0.59971 | 0.0063844 | 0.59971 | 0.75908 | 0.75908 | 1.2604e+17 | 0.016685 | 8.4243e+14 | 8.4243e+14 | 8.4243e+14 | -2.7202e+17 |  |  | 8.4243e+14 | 0.00040646 | 0.00040646 | 2460.2601 | 2459.0821 | 2459.0821 | 2519.1634 |  |  | 0.59364 | 0.87376 | 0.0026229 | 0.0044183
0.0072499 | 0.58076 | 0.0063072 | 0.58076 | 0.74991 | 0.74991 | 1.2452e+17 | 0.016155 | 8.1535e+14 | 8.1535e+14 | 8.1535e+14 | -2.6382e+17 |  |  | 8.1535e+14 | 0.00039917 | 0.00039917 | 2505.2238 | 2504.1203 | 2504.1203 | 2480.0912 |  |  | 0.59206 | 0.86113 | 0.0028777 | 0.0048604
0.0073499 | 0.56291 | 0.0059962 | 0.56291 | 0.71293 | 0.71293 | 1.1838e+17 | 0.015657 | 7.8989e+14 | 7.8989e+14 | 7.8989e+14 | -2.5713e+17 |  |  | 7.8989e+14 | 0.00040544 | 0.00040544 | 2466.4591 | 2465.4234 | 2465.4234 | 2443.3508 |  |  | 0.59053 | 0.82133 | 0.0036929 | 0.0062535
0.0074499 | 0.54514 | 0.0053942 | 0.54514 | 0.64135 | 0.64135 | 1.0649e+17 | 0.015161 | 7.6454e+14 | 7.6454e+14 | 7.6454e+14 | -2.4927e+17 |  |  | 7.6454e+14 | 0.00043181 | 0.00043181 | 2315.8504 | 2314.8801 | 2314.8801 | 2406.7888 |  |  | 0.58896 | 0.74644 | 0.0052268 | 0.0088748
0.0075499 | 0.52739 | 0.0055387 | 0.52739 | 0.65853 | 0.65853 | 1.0934e+17 | 0.014665 | 7.3927e+14 | 7.3927e+14 | 7.3927e+14 | -2.4376e+17 |  |  | 7.3927e+14 | 0.00040938 | 0.00040938 | 2442.7117 | 2441.8045 | 2441.8045 | 2370.3206 |  |  | 0.58734 | 0.7613 | 0.004907 | 0.0083546
0.0076499 | 0.51106 | 0.0051378 | 0.51106 | 0.61087 | 0.61087 | 1.0143e+17 | 0.01421 | 7.1602e+14 | 7.1602e+14 | 7.1602e+14 | -2.3595e+17 |  |  | 7.1602e+14 | 0.00042495 | 0.00042495 | 2353.2088 | 2352.3578 | 2352.3578 | 2336.778 |  |  | 0.58581 | 0.71034 | 0.005939 | 0.010138
0.0077499 | 0.49435 | 0.0046993 | 0.49435 | 0.55874 | 0.55874 | 92773848561495680 | 0.013743 | 6.9227e+14 |  | 6.9227e+14 | -2.2886e+17 |  |  | 6.9227e+14 | 0.00044542 |  |  | 2244.2903 |  | 2302.5036 |  |  | 0.58419 | 0.65522 | 0.0070497 | 0.012067
0.0078499 | 0.47872 | 0.0048077 | 0.47872 | 0.57162 | 0.57162 | 94913873880863776 | 0.013307 | 6.7007e+14 |  | 6.7007e+14 | -2.2492e+17 |  |  | 6.7007e+14 | 0.00042387 |  |  | 2358.4543 |  | 2270.482 |  |  | 0.58263 | 0.66645 | 0.0068019 | 0.011674
0.0079499 | 0.46304 | 0.0041442 | 0.46304 | 0.49273 | 0.49273 | 81813523314200624 | 0.01287 | 6.4782e+14 |  | 6.4782e+14 | -2.182e+17 |  |  | 6.4782e+14 | 0.00046707 |  |  | 2140.29 |  | 2238.368 |  |  | 0.58102 | 0.58471 | 0.0084452 | 0.014535
0.0080499 | 0.4475 | 0.0044034 | 0.4475 | 0.52355 | 0.52355 | 86932211181412432 | 0.012437 | 6.2579e+14 |  | 6.2579e+14 | -2.1118e+17 |  |  | 6.2579e+14 | 0.00043067 |  |  | 2321.3355 |  | 2206.5832 |  |  | 0.57938 | 0.61258 | 0.0078562 | 0.01356
0.0081499 | 0.43305 | 0.0041427 | 0.43305 | 0.49255 | 0.49255 | 81784700094954528 | 0.012034 | 605310548051345 |  | 605310548051345 | -2.0415e+17 |  |  | 605310548051345 | 0.00044103 |  |  | 2266.8278 |  | 2177.0412 |  |  | 0.5778 | 0.57862 | 0.0085216 | 0.014748
0.0082499 | 0.41922 | 0.0036393 | 0.41922 | 0.4327 | 0.4327 | 71846217195726704 | 0.011648 | 5.8574e+14 |  | 5.8574e+14 | -1.9672e+17 |  |  | 5.8574e+14 | 0.0004789 |  |  | 2087.552 |  | 2148.7997 |  |  | 0.57624 | 0.51563 | 0.009769 | 0.016953
0.0083499 | 0.4051 | 0.0039156 | 0.4051 | 0.46556 | 0.46556 | 77302294663565856 | 0.011255 | 5.6576e+14 |  | 5.6576e+14 | -1.9395e+17 |  |  | 5.6576e+14 | 0.00043578 |  |  | 2294.1808 |  | 2119.9828 |  |  | 0.5746 | 0.54732 | 0.0091039 | 0.015844
0.0084499 | 0.39197 | 0.003545 | 0.39197 | 0.42149 | 0.42149 | 69985737619183816 | 0.010889 | 5.472e+14 |  | 5.472e+14 | -1.8896e+17 |  |  | 5.472e+14 | 0.00046031 |  |  | 2171.9328 |  | 2093.1992 |  |  | 0.57303 | 0.50115 | 0.010005 | 0.01746
0.0085499 | 0.37809 | 0.0032135 | 0.37809 | 0.38207 | 0.38207 | 63440497818599840 | 0.010502 | 5.276e+14 |  | 5.276e+14 | -1.8327e+17 |  |  | 5.276e+14 | 0.00048423 |  |  | 2064.656 |  | 2064.9246 |  |  | 0.57132 | 0.45934 | 0.010811 | 0.018923
0.0086499 | 0.36585 | 0.003461 | 0.36585 | 0.41151 | 0.41151 | 68327810254603232 | 0.010162 | 5.1034e+14 |  | 5.1034e+14 | -1.7908e+17 |  |  | 5.1034e+14 | 0.00044178 |  |  | 2263.1632 |  | 2040.0104 |  |  | 0.56976 | 0.48701 | 0.01023 | 0.017955
0.0087499 | 0.35324 | 0.0029674 | 0.35324 | 0.35281 | 0.35281 | 58582008601842544 | 0.0098106 | 492560090053309 |  | 492560090053309 | -1.7196e+17 |  |  | 492560090053309 | 0.00048824 |  |  | 2047.7666 |  | 2014.3615 |  |  | 0.5681 | 0.42531 | 0.011427 | 0.020114
0.0088499 | 0.34091 | 0.0030821 | 0.34091 | 0.36646 | 0.36646 | 60847395183000208 | 0.0094672 | 475180905095689 |  | 475180905095689 | -1.6751e+17 |  |  | 475180905095689 | 0.00045833 |  |  | 2181.4382 |  | 1989.2863 |  |  | 0.56643 | 0.43707 | 0.01116 | 0.019702
0.0089499 | 0.32956 | 0.0030553 | 0.32956 | 0.36326 | 0.36326 | 60317126916595888 | 0.0091513 | 4.592e+14 |  | 4.592e+14 | -1.6171e+17 |  |  | 4.592e+14 | 0.00044871 |  |  | 2228.2669 |  | 1966.2317 |  |  | 0.56484 | 0.43144 | 0.01124 | 0.0199
0.0090499 | 0.31823 | 0.0023795 | 0.31823 | 0.28292 | 0.28292 | 46975924791848848 | 0.0088361 | 443262186820079 |  | 443262186820079 | -1.5614e+17 |  |  | 443262186820079 | 0.00053585 |  |  | 1865.8847 |  | 1943.2331 |  |  | 0.5632 | 0.34874 | 0.012838 | 0.022794
0.0091499 | 0.30714 | 0.0029479 | 0.30714 | 0.35049 | 0.35049 | 58196448689598456 | 0.0085274 | 4.2766e+14 |  | 4.2766e+14 | -1.519e+17 |  |  | 4.2766e+14 | 0.00043494 |  |  | 2298.8904 |  | 1920.7264 |  |  | 0.56155 | 0.41453 | 0.011507 | 0.020491
0.0092499 | 0.29662 | 0.0027189 | 0.29662 | 0.32326 | 0.32326 | 53675546492778368 | 0.0082348 | 4.1288e+14 |  | 4.1288e+14 | -1.4878e+17 |  |  | 4.1288e+14 | 0.00045096 |  |  | 2217.2228 |  | 1899.4001 |  |  | 0.55993 | 0.38598 | 0.012033 | 0.021491
0.0093499 | 0.28634 | 0.0023175 | 0.28634 | 0.27554 | 0.27554 | 45751727651129128 | 0.0079487 | 3.9844e+14 |  | 3.9844e+14 | -1.4476e+17 |  |  | 3.9844e+14 | 0.00049908 |  |  | 2003.4426 |  | 1878.5631 |  |  | 0.5583 | 0.33657 | 0.012964 | 0.02322
0.0094499 | 0.27569 | 0.0026656 | 0.27569 | 0.31694 | 0.31694 | 52624880925464984 | 0.0076523 | 3.8349e+14 |  | 3.8349e+14 | -1.3973e+17 |  |  | 3.8349e+14 | 0.00043015 |  |  | 2324.5139 |  | 1856.9859 |  |  | 0.55655 | 0.37584 | 0.012158 | 0.021846
0.0095499 | 0.26606 | 0.0021954 | 0.26606 | 0.26103 | 0.26103 | 43341830135121152 | 0.0073846 | 3.6998e+14 |  | 3.6998e+14 | -1.3544e+17 |  |  | 3.6998e+14 | 0.00049029 |  |  | 2039.368 |  | 1837.5002 |  |  | 0.55492 | 0.31813 | 0.013243 | 0.023866
0.0096499 | 0.25698 | 0.0024352 | 0.25698 | 0.28953 | 0.28953 | 48074958090088896 | 0.007132 | 3.5725e+14 |  | 3.5725e+14 | -1.2882e+17 |  |  | 3.5725e+14 | 0.00043802 |  |  | 2282.8077 |  | 1819.13 |  |  | 0.55333 | 0.34384 | 0.012708 | 0.022966
0.0097499 | 0.24741 | 0.0022862 | 0.24741 | 0.27183 | 0.27183 | 45134792307676616 | 0.0068659 | 3.4384e+14 |  | 3.4384e+14 | -1.2537e+17 |  |  | 3.4384e+14 | 0.00044645 |  |  | 2239.6779 |  | 1799.7815 |  |  | 0.5516 | 0.32468 | 0.013038 | 0.023636
0.0098499 | 0.23906 | 0.0018121 | 0.23906 | 0.21546 | 0.21546 | 35774747986469896 | 0.0066338 | 3.3215e+14 |  | 3.3215e+14 | -1.2183e+17 |  |  | 3.3215e+14 | 0.0005248 |  |  | 1905.303 |  | 1782.9125 |  |  | 0.55004 | 0.26682 | 0.014115 | 0.025661
0.0099499 | 0.23027 | 0.0022023 | 0.23027 | 0.26184 | 0.26184 | 43476864943096040 | 0.0063895 | 3.1985e+14 |  | 3.1985e+14 | -1.1652e+17 |  |  | 3.1985e+14 | 0.00043362 |  |  | 2306.0072 |  | 1765.1638 |  |  | 0.54835 | 0.31097 | 0.013224 | 0.024116
0.01005 | 0.22203 | 0.0019464 | 0.22203 | 0.23142 | 0.23142 | 38425694479871648 | 0.0061607 | 3.0833e+14 |  | 3.0833e+14 | -1.1433e+17 |  |  | 3.0833e+14 | 0.00046486 |  |  | 2151.0298 |  | 1748.5465 |  |  | 0.54671 | 0.27962 | 0.013784 | 0.025213
0.01015 | 0.21428 | 0.0017457 | 0.21428 | 0.20756 | 0.20756 | 34463686349392068 | 0.0059453 | 2.9749e+14 |  | 2.9749e+14 | -1.0987e+17 |  |  | 2.9749e+14 | 0.00049399 |  |  | 2024.1679 |  | 1732.9036 |  |  | 0.54512 | 0.25388 | 0.014235 | 0.026114
0.01025 | 0.20617 | 0.0019884 | 0.20617 | 0.23641 | 0.23641 | 39254658162161928 | 0.00572 | 2.8616e+14 |  | 2.8616e+14 | -1.0502e+17 |  |  | 2.8616e+14 | 0.00042979 |  |  | 2326.5579 |  | 1716.5606 |  |  | 0.5434 | 0.28069 | 0.013681 | 0.025176
0.01035 | 0.1989 | 0.0015455 | 0.1989 | 0.18376 | 0.18376 | 30511549184407736 | 0.0055179 | 2.76e+14 |  | 2.76e+14 | -1.0001e+17 |  |  | 2.76e+14 | 0.00051502 |  |  | 1941.5455 |  | 1701.8962 |  |  | 0.5418 | 0.22592 | 0.014679 | 0.027093
0.01045 | 0.19214 | 0.0017936 | 0.19214 | 0.21325 | 0.21325 | 35408337747286576 | 0.0053302 | 2.6656e+14 |  | 2.6656e+14 | -94107883455966432 |  |  | 2.6656e+14 | 0.00044431 |  |  | 2250.5513 |  | 1688.2837 |  |  | 0.54028 | 0.25292 | 0.014127 | 0.026148
0.01055 | 0.18549 | 0.0018048 | 0.18549 | 0.21459 | 0.21459 | 35630237051619580 | 0.0051454 | 2.5728e+14 |  | 2.5728e+14 | -91381724692833776 |  |  | 2.5728e+14 | 0.00042852 |  |  | 2333.4894 |  | 1674.8904 |  |  | 0.53873 | 0.25311 | 0.014083 | 0.026141
0.01065 | 0.17917 | 0.0013805 | 0.17917 | 0.16413 | 0.16413 | 27253340893738800 | 0.0049699 | 2.4846e+14 |  | 2.4846e+14 | -89576019228833952 |  |  | 2.4846e+14 | 0.00051881 |  |  | 1927.3762 |  | 1662.1727 |  |  | 0.53721 | 0.2019 | 0.015006 | 0.027934
0.01075 | 0.17264 | 0.0018107 | 0.17264 | 0.21528 | 0.21528 | 35745727347913888 | 0.0047885 | 2.3935e+14 |  | 2.3935e+14 | -85537715743429312 |  |  | 2.3935e+14 | 0.00040147 |  |  | 2490.7433 |  | 1649.0276 |  |  | 0.53559 | 0.25134 | 0.014034 | 0.026203
0.01085 | 0.16646 | 0.001628 | 0.16646 | 0.19357 | 0.19357 | 32140653329742208 | 0.0046168 | 2.3074e+14 |  | 2.3074e+14 | -82511331278344160 |  |  | 2.3074e+14 | 0.00042598 |  |  | 2347.4603 |  | 1636.5937 |  |  | 0.53401 | 0.22835 | 0.014422 | 0.027008
0.01095 | 0.16118 | 0.0013024 | 0.16118 | 0.15485 | 0.15485 | 25710903825452544 | 0.0044701 | 2.2337e+14 |  | 2.2337e+14 | -78508293148876864 |  |  | 2.2337e+14 | 0.00050106 |  |  | 1995.7048 |  | 1625.9717 |  |  | 0.53261 | 0.18794 | 0.015138 | 0.028422
0.01105 | 0.15531 | 0.0017008 | 0.15531 | 0.20222 | 0.20222 | 33576878809299676 | 0.0043072 | 2.152e+14 |  | 2.152e+14 | -75813208761114144 |  |  | 2.152e+14 | 0.00038741 |  |  | 2581.1833 |  | 1614.1758 |  |  | 0.53102 | 0.23418 | 0.014233 | 0.026804
0.01115 | 0.15004 | 0.001338 | 0.15004 | 0.15908 | 0.15908 | 26414703665263180 | 0.004161 | 2.0787e+14 |  | 2.0787e+14 | -71210619446507656 |  |  | 2.0787e+14 | 0.00046341 |  |  | 2157.852 |  | 1603.5995 |  |  | 0.52954 | 0.1891 | 0.015029 | 0.028381
0.01125 | 0.14481 | 0.0015646 | 0.14481 | 0.18602 | 0.18602 | 30887435550466488 | 0.0040157 | 2.0058e+14 |  | 2.0058e+14 | -65533772606782688 |  |  | 2.0058e+14 | 0.00039579 |  |  | 2526.494 |  | 1593.087 |  |  | 0.52803 | 0.21365 | 0.014533 | 0.027522
0.01135 | 0.14085 | 0.0015611 | 0.14085 | 0.18561 | 0.18561 | 30819918146479044 | 0.0039059 | 1.9508e+14 |  | 1.9508e+14 | -63906694015229200 |  |  | 1.9508e+14 | 0.00038691 |  |  | 2584.5158 |  | 1585.1437 |  |  | 0.52686 | 0.21256 | 0.014521 | 0.027561
0.01145 | 0.13636 | 0.00115 | 0.13636 | 0.13673 | 0.13673 | 22703418058362712 | 0.0037812 | 1.8883e+14 |  | 1.8883e+14 | -60914428653994528 |  |  | 1.8883e+14 | 0.00049014 |  |  | 2040.1611 |  | 1576.1266 |  |  | 0.52549 | 0.16241 | 0.015405 | 0.029316
0.01155 | 0.1313 | 0.0016178 | 0.1313 | 0.19235 | 0.19235 | 31938101117779872 | 0.0036408 | 1.8179e+14 |  | 1.8179e+14 | -59538858807681280 |  |  | 1.8179e+14 | 0.00035245 |  |  | 2837.2359 |  | 1565.9765 |  |  | 0.52391 | 0.21745 | 0.014349 | 0.027389
0.01165 | 0.12769 | 0.0013932 | 0.12769 | 0.16564 | 0.16564 | 27503865998008000 | 0.0035405 | 1.7677e+14 |  | 1.7677e+14 | -56263675363309040 |  |  | 1.7677e+14 | 0.00039354 |  |  | 2541.0091 |  | 1558.7261 |  |  | 0.52274 | 0.18936 | 0.014831 | 0.028372
0.01175 | 0.12378 | 0.0012608 | 0.12378 | 0.14991 | 0.14991 | 24891613689347600 | 0.0034321 | 1.7134e+14 |  | 1.7134e+14 | -53493015601472352 |  |  | 1.7134e+14 | 0.00041883 |  |  | 2387.566 |  | 1550.8918 |  |  | 0.52145 | 0.17246 | 0.015103 | 0.028964
0.01185 | 0.11989 | 0.0015377 | 0.11989 | 0.18283 | 0.18283 | 30357364703372084 | 0.003324 | 1.6592e+14 |  | 1.6592e+14 | -50846834328200136 |  |  | 1.6592e+14 | 0.00034245 |  |  | 2920.129 |  | 1543.0795 |  |  | 0.52013 | 0.20426 | 0.014486 | 0.027851
0.01195 | 0.11596 | 0.0011818 | 0.11596 | 0.14051 | 0.14051 | 23330026948000572 | 0.0032151 | 1.6047e+14 |  | 1.6047e+14 | -48408996430936616 |  |  | 1.6047e+14 | 0.00042041 |  |  | 2378.5953 |  | 1535.2091 |  |  | 0.51876 | 0.16091 | 0.015235 | 0.029368
0.01205 | 0.1133 | 0.0014498 | 0.1133 | 0.17238 | 0.17238 | 28622048969308828 | 0.0031414 | 1.5678e+14 |  | 1.5678e+14 | -45251984094610592 |  |  | 1.5678e+14 | 0.00034522 |  |  | 2896.6488 |  | 1529.8858 |  |  | 0.51781 | 0.19146 | 0.014654 | 0.028299
0.01215 | 0.10965 | 0.0014073 | 0.10965 | 0.16733 | 0.16733 | 27783411740833208 | 0.00304 | 1.517e+14 |  | 1.517e+14 | -43436689078856712 |  |  | 1.517e+14 | 0.00034452 |  |  | 2902.5858 |  | 1522.5657 |  |  | 0.51647 | 0.18564 | 0.014721 | 0.028503
0.01225 | 0.10675 | 0.0011471 | 0.10675 | 0.13638 | 0.13638 | 22645574200560604 | 0.0029596 | 147678237239640 |  | 147678237239640 | -42899658417462936 |  |  | 147678237239640 | 0.00040304 |  |  | 2481.0867 |  | 1516.7558 |  |  | 0.51539 | 0.15447 | 0.015252 | 0.029594
0.01235 | 0.1036 | 0.0015836 | 0.1036 | 0.18829 | 0.18829 | 31263321916525240 | 0.002872 | 1.433e+14 |  | 1.433e+14 |  |  |  | 1.433e+14 | [5 empty] | 1510.4382 |  |  | 0.51417
0.01245 | 0.10088 | 0.0013619 | 0.10088 | 0.16193 | 0.16193 | 26886930654555480 | 0.0027967 | 1.3953e+14 |  | 1.3953e+14 |  |  |  | 1.3953e+14 | [5 empty] | 1505.001 |  |  | 0.51311
Flag marking autoexit | true
Flag for automatic flash | false
Gain last used on channel 3 | 5
Gain last used on channel 2 | 5
Gain last used on channel 1 | 10
Gain last used on channel 0 | 20
Maximum value measured on channel 3 | 1
Maximum value measured on channel 2 | 1
Maximum value measured on channel 1 | 5.2886
Maximum value measured on channel 0 | 0.44356
Number of points to transfer to Excel | 125
Pulse length in us | 12500
Data frequency in us | 0.2
Frequency of AD-conversions | 0.02
Timeout in ms | 1000
Trigger level in volts | 0.05
Dark voltage average value | -0.8454
Flag for dark voltage subtraction | true
Flag for refcell normalization | true
Flag for photoconductance normalization | false
Time in us for measuring baseline mean | 20000
Time in s between flashes | 6
Gain used for baseline measurement on ch 1 | 1
Gain used for thermocouple measurement on ch 0 | 20
Thickness | 1
Do autoscale gains | false
[Analysis parameters]
Use lifetime limits | 1
Low lifetime limit | 1
High lifetime limit | 1
Use resistivity limits | 1
Low resistivity limit | 1
High resistivity limit | 1
Use trapping limits | 1
Low trapping limit | 1
High trapping limit | 1
Pass wafer on timeout | false
TTL LOW width | 1
TTL HIGH width | 1
Transmission fraction | 1
Light bias | 1
Wafer width | 1
MCD to evaluate tau | 1
Fit range | 1
User-input resistivity | 1
Analysis type (0:QSS, 1:Trans, 2:Gen) | 1
Dopant type (0:p, 1:n) | 1
Use fixed density (0:yes, 1:no) | 1
Percent of max N | 1
Coupling_channel_0 | 1
Coupling_channel_1 | 1
Use_linear_divisions | 1
Auger_coef_used | 1
Plot_type_left_graph | 1
Plot_type_right_graph | 1
Right_graph_x_min | 1
Right_graph_x_max | 1
Right_graph_left_y_min | 1
Right_graph_left_y_max | 1
Right_graph_right_y_min | 1
Right_graph_right_y_max | 1
Left_graph_x_min | 1
Left_graph_x_max | 1
Left_graph_left_y_min | 1
Left_graph_left_y_max | 1
Left_graph_right_y_min | 1
Left_graph_right_y_max | 1
[Communications]
Use database | false
Database location | 1.0default
Semaphore location | 1.0default
Autosave full data | false
Autosave (full data) location | 1.0default
Autosave stats | false
Autosave (stats) location | 1.0default
Autosave id (0:date, 1:lot) | 1
File suffix | 1
File suffix for lotIDs | 1
Current date string | 1.0default
Current lot ID | 1.0default
Current lot start index | 1
Current lot end index | 1
Current recipe name | 1.0default
Default_recipe_path | 1.0default
[Instrument constants]
Refcell constant | 1
Cal const A | 1
Cal const B | 1
Cal const C | 1
Area fraction | 1
Primary refcell constant | 1
Secondary refcell constant | 1
Index_of_flashmode | 1
Cal const B | 1
Cal const C | 1
Area fraction | 1
Primary refcell constant | 1
Secondary refcell constant | 1

## Settings
Instrument Calibration |  | Summary Log Settings
Instument Type | WCT-120 | When user clicks on Summary button, send the results to the following files |  | Baseline Air Voltage Measurements
Serial Number | WCT-0318 | TRUE | C:\SintonInstruments\LifetimeSummaryData\summary.xlsm | -0.88585
Ref cell (V/sun) | 0.0084106 | FALSE | C:\SintonInstruments\LifetimeSummaryData\summary.txt | -0.88594
A | 0.00019858 | FALSE | C:\SintonInstruments\LifetimeSummaryData\summary.accdb | -0.88585
B | 0.027687 | TRUE | automatically Log results to specified summary files after each measurement | -0.88581
Offset = Air-C | -7.1842e-05 |  |  | -0.88572
C | -0.88568 |  |  | -0.88577
Air Voltage | -0.88575 | Autosave Settings |  | -0.88558
|  | TRUE | Automatically save file after each measurement | -0.88557
Analysis Parameters\Settings |  |  |  | -0.8857
Auger Coefficient | 1.66e-30 | Rescale Setting |  | -0.8857
Coupling Mode | dc | TRUE | Automatically Rescale Graphs after each measurement | Air Voltage Stdev
Illuminated Area Fraction | 1 |  |  | 0.00011892
Fit Range | 30% | Show Taking Data Window Setting |  | Date and Time of Last "Zero Instrument"
Auger Model | Sinton | TRUE | Show the Taking Data window during measurement | 2022-07-07 15:23
Bandgap Narrowing Analysis | OFF
*BGN from Achim Kimmerle, Fraunhofer ISE |  | Averaging Settings
|  | 3 | Number of Flashes to average (initial value)
Acceptance Parameters |  | 6 | Delay between flashes QSS/Generalized (1/1) (sec)
Enable Pass/Fail Message | FALSE | 0.3 | Delay between flashes Transient/Generalized (1/64) (sec)
Minimum Lifetime (us) | 1
Maximum Lifetime (us) | 20 | Baseline Voltage Range
Pass or Fail |  | 1 | Air Voltage Input Range (Volts +/-)
Data Acquisition Time Scale Settings
|  | 12500 | QSS/Generalized Mode Acquisition Time (µs)
|  | 100000 | Transient Mode Total Acquisition Time (µs)
|  | 3 | Transient Mode # of Taus Beyond 150mV
Note on B3: Instrument Type
Note on B4: Instrument Serial Number
Note on F4: Save summary of measurement of to Excel summary File. Location can be specified as file path, otherwise default is to save to same location as measurement file.
Note on B5: Calibration constant of the reference cell that measures the light intensity. Do Not Change this Value.
Note on F5: Save summary of measurement of to Text File. Location can be specified as file path, otherwise default is to save to same location as measurement file.
Note on B6: Calibration of the photoconductivity sensor.
Note on F6: Save summary of measurement of to Access Database. Location can be specified as file path, otherwise default is to save to same location as measurement file.
Note on F7: True to automatically save summary of each measurement to files specified above.
Note on F11: True to Automatically save a new Excel measurement file as SampleName.xlsm.
Note on B13: Don't change this value
Note on B14: Normally Set do DC.
Note on F14: True to automatically rescale graphs after each measurement.
Note on B15: For standard instrument = 1.
Note on B16: Fit will cover specified carrier density +/- Fit Range.
Note on B17: The Auger analysis can be selected here. The Options are "Sinton" or "Richter." When Richter analysis is selected, bandgap narrowing analysis is turned on.
Note on F17: Set to TRUE to enable the "Taking Data" Screen. False to just take a measurement without making adjustments.
Note on B18: When the Richter Auger model is used, bandgap narrowing analysis is used in the analysis of the data.
Note on F20: Initial Value for Average Function in "Taking Data" Window.
Note on F21: Time between consectutive flashes in Average mode for QSS measurements. Must be set to 6 seconds for flash to recharge fully.
Note on F22: Time between consectutive flashes in Average mode for Transient measurements. Must be set to 0.3 seconds for flash to recharge fully at 1/64.
Note on F25: +/- voltage range for the "Zero Instrument" routine. Typically set to 1 volt.

## Summary
Sample Name | Measurement Date | Measurement Time | Thickness (cm) | Wafer Doping (cm-3) | Optical Constant | Specified MCD | Light Bias (Suns) | Analysis Mode | Lifetime at Spec. MCD | Jo (A/cm2) 25 C | Tau fit @ -NA | Ohms/sq | Ohm-cm | Max Suns | Traps | Max MCD | 1 Sun Implied Voc | Implied FF | Ref Cell Cal Const | Calibration A | Calibration B | Calibration C | Instrument Serial Number
nSi_AlOx10B_450 | 2022-07-07 | 16:55 | 0.07 | 1.1092e+14 | 0.7 | 1000000000000000 | 0 | Generalized (1/1) | 363.872 | 1.1968e-12 | 874.4002 | 896.4372 | 62.7506 | 48.8362 | 0 | 8318371307832016 | 0.59746 | 73.6699 | 0.0084106 | 0.00019858 | 0.027687 | -0.88568 | WCT-0318
The row above (row 2) is copied to summary.xlsm when the user presses the "log data to summary" button.
Change row 2 to control what gets sent to the summary file
The headings in row 1 are for reference and have no effect on program operation
Row 10 is for updating an access database File Summary.mdb The names below must match the column headings in the db file.
For both text files and database files only that data that has a heading below gets updates (see manual for more details)
Name | Date | Time | Width | Doping | OptConst | SpecMCD | BiasLight | Analysis | Lifetime | J0 | TauFit | Ohmsqr | Ohmcm | MaxSuns | Traps | MaxMCD | 1SunImpVoc | ImpFF | RefCellCal | Acal | Bcal | Ccal | InstrumentSN

## RawData
Time (s) | Photovoltage | Reference Voltage | Apparent CD | Tau (sec) | 1/Tau corrected | Minority Carrier Density | Implied Voc | Implied Suns
4.99e-05 | 1.2987 | 0.2308 | 1.8605e+15 | 0 | 0 | 1.8605e+15
0.0001499 | 3.5962 | 0.34791 | 5467332468520812 | 0 | 0 | 5467332468520812
0.0002499 | 4.7733 | 0.41074 | 7470796002349974 | 7.4624e-05 | 13307.873 | 7470796002349974
0.0003499 | 5.2317 | 0.40933 | 8280128858683653 | 7.5737e-05 | 13089.7263 | 8280128858683653
0.0004499 | 5.2532 | 0.3818 | 8318371307832016 | 7.7625e-05 | 12767.5189 | 8318371307832016
0.0005499 | 5.1309 | 0.35719 | 8100663770522564 | 7.8848e-05 | 12573.6537 | 8100663770522564 | 0.7071 | 43.3119
0.0006499 | 4.9721 | 0.3325 | 7819653846977320 | 8.1032e-05 | 12239.2615 | 7819653846977320 | 0.7053 | 40.6826
0.0007499 | 4.8067 | 0.30996 | 7529106303172681 | 8.3395e-05 | 11896.9994 | 7529106303172681 | 0.70337 | 38.0611
0.0008499 | 4.6414 | 0.28866 | 7240838114708719 | 8.5945e-05 | 11548.3387 | 7240838114708719 | 0.70138 | 35.518
0.0009499 | 4.478 | 0.26829 | 6957994349400993 | 8.8714e-05 | 11191.754 | 6957994349400993 | 0.69935 | 33.065
0.0010499 | 4.3196 | 0.2496 | 6685751356152515 | 9.1505e-05 | 10854.2131 | 6685751356152515 | 0.69731 | 30.8025
0.0011499 | 4.1662 | 0.23242 | 6423987082665524 | 9.4322e-05 | 10533.5193 | 6423987082665524 | 0.69528 | 28.7126
0.0012499 | 4.0182 | 0.216 | 6173260391428934 | 9.742e-05 | 10201.6039 | 6173260391428934 | 0.69325 | 26.7145
0.0013499 | 3.8764 | 0.20158 | 5934470342486079 | 0.00010026 | 9915.3577 | 5934470342486079 | 0.69124 | 24.953
0.0014499 | 3.7398 | 0.18764 | 5705928596947935 | 0.00010345 | 9612.2565 | 5705928596947935 | 0.68924 | 23.2523
0.0015499 | 3.6087 | 0.1746 | 5487981409533569 | 0.00010681 | 9312.2718 | 5487981409533569 | 0.68726 | 21.6607
0.0016499 | 3.4823 | 0.16314 | 5279140938784232 | 0.00010986 | 9056.0371 | 5279140938784232 | 0.68529 | 20.2578
0.0017499 | 3.3606 | 0.15196 | 5079038322726614 | 0.00011335 | 8779.278 | 5079038322726614 | 0.68332 | 18.89
0.0018499 | 3.2443 | 0.14189 | 4888945383497467 | 0.00011674 | 8526.1375 | 4888945383497467 | 0.68138 | 17.6548
0.0019499 | 3.1324 | 0.13259 | 4707078034554299 | 0.00012018 | 8284.3576 | 4707078034554299 | 0.67946 | 16.5125
0.0020499 | 3.0256 | 0.12359 | 4534386484947477 | 0.00012406 | 8026.2452 | 4534386484947477 | 0.67756 | 15.4082
0.0021499 | 2.9233 | 0.11597 | 4369637826078056 | 0.0001273 | 7823.657 | 4369637826078056 | 0.67568 | 14.4707
0.0022499 | 2.8243 | 0.10834 | 4211130389768981 | 0.00013118 | 7593.9117 | 4211130389768981 | 0.67381 | 13.5339
0.0023499 | 2.7294 | 0.10114 | 4059693690977409 | 0.0001353 | 7363.8319 | 4059693690977409 | 0.67195 | 12.6499
0.0024499 | 2.6383 | 0.095122 | 3.9151e+15 | 0.00013863 | 7188.0199 | 3.9151e+15 | 0.67011 | 11.9058
0.0025499 | 2.5512 | 0.088833 | 3.7774e+15 | 0.00014305 | 6966.8708 | 3.7774e+15 | 0.6683 | 11.1322
0.0026499 | 2.4673 | 0.083323 | 3645283068023184 | 0.00014702 | 6779.8278 | 3645283068023184 | 0.6665 | 10.453
0.0027499 | 2.3862 | 0.078244 | 3.5182e+15 | 0.00015095 | 6604.3517 | 3.5182e+15 | 0.6647 | 9.8259
0.0028499 | 2.3087 | 0.073085 | 3397013458479484 | 0.00015583 | 6398.2327 | 3397013458479484 | 0.66293 | 9.1904
0.0029499 | 2.2343 | 0.069026 | 3.2812e+15 | 0.00015923 | 6262.3803 | 3.2812e+15 | 0.66117 | 8.6875
0.0030499 | 2.1624 | 0.0647 | 3169704805721946 | 0.00016387 | 6085.7113 | 3169704805721946 | 0.65943 | 8.1545
0.0031499 | 2.0926 | 0.060496 | 3.0619e+15 | 0.00016903 | 5900.4519 | 3.0619e+15 | 0.65768 | 7.6365
0.0032499 | 2.0258 | 0.057276 | 2958950718217559 | 0.00017241 | 5785.5742 | 2958950718217559 | 0.65595 | 7.2352
0.0033499 | 1.9616 | 0.053619 | 2.8604e+15 | 0.00017776 | 5611.8404 | 2.8604e+15 | 0.65425 | 6.7836
0.0034499 | 1.8999 | 0.050447 | 2.7659e+15 | 0.00018248 | 5467.4266 | 2.7659e+15 | 0.65255 | 6.3902
0.0035499 | 1.8396 | 0.047659 | 2.6739e+15 | 0.0001865 | 5349.9802 | 2.6739e+15 | 0.65085 | 6.0441
0.0036499 | 1.7826 | 0.044465 | 2587131206289060 | 0.00019305 | 5168.7871 | 2587131206289060 | 0.64919 | 5.6496
0.0037499 | 1.7265 | 0.042314 | 2.502e+15 | 0.00019608 | 5089.6146 | 2.502e+15 | 0.6475 | 5.3794
0.0038499 | 1.6733 | 0.0397 | 2421528536526037 | 0.00020188 | 4943.76 | 2421528536526037 | 0.64586 | 5.0569
0.0039499 | 1.6216 | 0.037158 | 2.3435e+15 | 0.00020834 | 4790.685 | 2.3435e+15 | 0.64421 | 4.742
0.0040499 | 1.571 | 0.03537 | 2267154283863490 | 0.00021151 | 4719.3277 | 2267154283863490 | 0.64255 | 4.5188
0.0041499 | 1.5228 | 0.03314 | 2194801256365321 | 0.00021809 | 4577.3234 | 2194801256365321 | 0.64093 | 4.2427
0.0042499 | 1.4756 | 0.03144 | 2.1241e+15 | 0.00022227 | 4491.4815 | 2.1241e+15 | 0.63928 | 4.0287
0.0043499 | 1.4302 | 0.029848 | 2.0562e+15 | 0.00022639 | 4410.093 | 2.0562e+15 | 0.63766 | 3.8291
0.0044499 | 1.387 | 0.027914 | 1.9917e+15 | 0.000234 | 4266.99 | 1.9917e+15 | 0.63606 | 3.5884
0.0045499 | 1.3445 | 0.02674 | 1.9286e+15 | 0.00023636 | 4224.5739 | 1.9286e+15 | 0.63445 | 3.4398
0.0046499 | 1.3039 | 0.025215 | 1868287424157004 | 0.00024234 | 4120.6201 | 1868287424157004 | 0.63286 | 3.2501
0.0047499 | 1.2639 | 0.023425 | 1.809e+15 | 0.00025181 | 3965.8443 | 1.809e+15 | 0.63125 | 3.0286
0.0048499 | 1.2257 | 0.022517 | 1.7525e+15 | 0.00025371 | 3936.4751 | 1.7525e+15 | 0.62967 | 2.912
0.0049499 | 1.1886 | 0.021107 | 1.6977e+15 | 0.00026153 | 3818.9123 | 1.6977e+15 | 0.62809 | 2.7367
0.0050499 | 1.1531 | 0.019729 | 1645338890713628 | 0.00027034 | 3694.5918 | 1645338890713628 | 0.62653 | 2.5658
0.0051499 | 1.1176 | 0.018978 | 1593108604489832 | 0.00027203 | 3671.9185 | 1593108604489832 | 0.62493 | 2.469
0.0052499 | 1.0839 | 0.017658 | 1.5437e+15 | 0.00028234 | 3537.833 | 1.5437e+15 | 0.62336 | 2.305
0.0053499 | 1.0517 | 0.017037 | 1.4965e+15 | 0.00028368 | 3521.4402 | 1.4965e+15 | 0.62182 | 2.224
0.0054499 | 1.0199 | 0.016144 | 1.4499e+15 | 0.0002893 | 3453.1247 | 1.4499e+15 | 0.62025 | 2.1129
0.0055499 | 0.98897 | 0.015 | 1.4048e+15 | 0.0003004 | 3325.6657 | 1.4048e+15 | 0.61868 | 1.9715
0.0056499 | 0.95838 | 0.01463 | 1.3601e+15 | 0.00029829 | 3349.3211 | 1.3601e+15 | 0.61709 | 1.9223
0.0057499 | 0.92967 | 0.0137 | 1.3183e+15 | 0.00030771 | 3246.9353 | 1.3183e+15 | 0.61554 | 1.8062
0.0058499 | 0.9008 | 0.012726 | 1.2764e+15 | 0.00031951 | 3127.1013 | 1.2764e+15 | 0.61395 | 1.6841
0.0059499 | 0.87327 | 0.01241 | 1.2364e+15 | 0.00031741 | 3147.9169 | 1.2364e+15 | 0.61238 | 1.6421
0.0060499 | 0.847 | 0.011465 | 1198299764662150 | 0.00033151 | 3014.159 | 1198299764662150 | 0.61084 | 1.5239
0.0061499 | 0.82034 | 0.010958 | 1159718392552213 | 0.0003352 | 2981.0456 | 1159718392552213 | 0.60923 | 1.4586
0.0062499 | 0.79569 | 0.01047 | 1.1241e+15 | 0.00033944 | 2943.8881 | 1.1241e+15 | 0.6077 | 1.3961
0.0063499 | 0.77131 | 0.0095895 | 1.0889e+15 | 0.00035685 | 2800.3475 | 1.0889e+15 | 0.60614 | 1.2864
0.0064499 | 0.74738 | 0.009548 | 1.0544e+15 | 0.00034796 | 2872.045 | 1.0544e+15 | 0.60456 | 1.2774
0.0065499 | 0.7243 | 0.0088995 | 1.0211e+15 | 0.00035995 | 2776.45 | 1.0211e+15 | 0.60299 | 1.196
0.0066499 | 0.70215 | 0.0082081 | 9.8927e+14 | 0.00037596 | 2658.2516 | 9.8927e+14 | 0.60145 | 1.1093
0.0067499 | 0.68034 | 0.0082701 | 9.5793e+14 | 0.00036296 | 2753.6113 | 9.5793e+14 | 0.59988 | 1.1126
0.0068499 | 0.6593 | 0.0076744 | 9.2775e+14 | 0.00037682 | 2652.3468 | 9.2775e+14 | 0.59832 | 1.0379
0.0069499 | 0.63888 | 0.0071275 | 8.9848e+14 | 0.00039086 | 2557.1347 | 8.9848e+14 | 0.59676 | 0.9691
0.0070499 | 0.61911 | 0.0070226 | 8.7018e+14 | 0.0003848 | 2597.522 | 8.7018e+14 | 0.59521 | 0.95336
0.0071499 | 0.59971 | 0.0063844 | 8.4243e+14 | 0.00040646 | 2459.0821 | 8.4243e+14 | 0.59364 | 0.87376
0.0072499 | 0.58076 | 0.0063072 | 8.1535e+14 | 0.00039917 | 2504.1203 | 8.1535e+14 | 0.59206 | 0.86113
0.0073499 | 0.56291 | 0.0059962 | 7.8989e+14 | 0.00040544 | 2465.4234 | 7.8989e+14 | 0.59053 | 0.82133
0.0074499 | 0.54514 | 0.0053942 | 7.6454e+14 | 0.00043181 | 2314.8801 | 7.6454e+14 | 0.58896 | 0.74644
0.0075499 | 0.52739 | 0.0055387 | 7.3927e+14 | 0.00040938 | 2441.8045 | 7.3927e+14 | 0.58734 | 0.7613
0.0076499 | 0.51106 | 0.0051378 | 7.1602e+14 | 0.00042495 | 2352.3578 | 7.1602e+14 | 0.58581 | 0.71034
0.0077499 | 0.49435 | 0.0046993 | 6.9227e+14 | 0.00044542 | 2244.2903 | 6.9227e+14 | 0.58419 | 0.65522
0.0078499 | 0.47872 | 0.0048077 | 6.7007e+14 | 0.00042387 | 2358.4543 | 6.7007e+14 | 0.58263 | 0.66645
0.0079499 | 0.46304 | 0.0041442 | 6.4782e+14 | 0.00046707 | 2140.29 | 6.4782e+14 | 0.58102 | 0.58471
0.0080499 | 0.4475 | 0.0044034 | 6.2579e+14 | 0.00043067 | 2321.3355 | 6.2579e+14 | 0.57938 | 0.61258
0.0081499 | 0.43305 | 0.0041427 | 605310548051345 | 0.00044103 | 2266.8278 | 605310548051345 | 0.5778 | 0.57862
0.0082499 | 0.41922 | 0.0036393 | 5.8574e+14 | 0.0004789 | 2087.552 | 5.8574e+14 | 0.57624 | 0.51563
0.0083499 | 0.4051 | 0.0039156 | 5.6576e+14 | 0.00043578 | 2294.1808 | 5.6576e+14 | 0.5746 | 0.54732
0.0084499 | 0.39197 | 0.003545 | 5.472e+14 | 0.00046031 | 2171.9328 | 5.472e+14 | 0.57303 | 0.50115
0.0085499 | 0.37809 | 0.0032135 | 5.276e+14 | 0.00048423 | 2064.656 | 5.276e+14 | 0.57132 | 0.45934
0.0086499 | 0.36585 | 0.003461 | 5.1034e+14 | 0.00044178 | 2263.1632 | 5.1034e+14 | 0.56976 | 0.48701
0.0087499 | 0.35324 | 0.0029674 | 492560090053309 | 0.00048824 | 2047.7666 | 492560090053309 | 0.5681 | 0.42531
0.0088499 | 0.34091 | 0.0030821 | 475180905095689 | 0.00045833 | 2181.4382 | 475180905095689 | 0.56643 | 0.43707
0.0089499 | 0.32956 | 0.0030553 | 4.592e+14 | 0.00044871 | 2228.2669 | 4.592e+14 | 0.56484 | 0.43144
0.0090499 | 0.31823 | 0.0023795 | 443262186820079 | 0.00053585 | 1865.8847 | 443262186820079 | 0.5632 | 0.34874
0.0091499 | 0.30714 | 0.0029479 | 4.2766e+14 | 0.00043494 | 2298.8904 | 4.2766e+14 | 0.56155 | 0.41453
0.0092499 | 0.29662 | 0.0027189 | 4.1288e+14 | 0.00045096 | 2217.2228 | 4.1288e+14 | 0.55993 | 0.38598
0.0093499 | 0.28634 | 0.0023175 | 3.9844e+14 | 0.00049908 | 2003.4426 | 3.9844e+14 | 0.5583 | 0.33657
0.0094499 | 0.27569 | 0.0026656 | 3.8349e+14 | 0.00043015 | 2324.5139 | 3.8349e+14 | 0.55655 | 0.37584
0.0095499 | 0.26606 | 0.0021954 | 3.6998e+14 | 0.00049029 | 2039.368 | 3.6998e+14 | 0.55492 | 0.31813
0.0096499 | 0.25698 | 0.0024352 | 3.5725e+14 | 0.00043802 | 2282.8077 | 3.5725e+14 | 0.55333 | 0.34384
0.0097499 | 0.24741 | 0.0022862 | 3.4384e+14 | 0.00044645 | 2239.6779 | 3.4384e+14 | 0.5516 | 0.32468
0.0098499 | 0.23906 | 0.0018121 | 3.3215e+14 | 0.0005248 | 1905.303 | 3.3215e+14 | 0.55004 | 0.26682
0.0099499 | 0.23027 | 0.0022023 | 3.1985e+14 | 0.00043362 | 2306.0072 | 3.1985e+14 | 0.54835 | 0.31097
0.01005 | 0.22203 | 0.0019464 | 3.0833e+14 | 0.00046486 | 2151.0298 | 3.0833e+14 | 0.54671 | 0.27962
0.01015 | 0.21428 | 0.0017457 | 2.9749e+14 | 0.00049399 | 2024.1679 | 2.9749e+14 | 0.54512 | 0.25388
0.01025 | 0.20617 | 0.0019884 | 2.8616e+14 | 0.00042979 | 2326.5579 | 2.8616e+14 | 0.5434 | 0.28069
0.01035 | 0.1989 | 0.0015455 | 2.76e+14 | 0.00051502 | 1941.5455 | 2.76e+14 | 0.5418 | 0.22592
0.01045 | 0.19214 | 0.0017936 | 2.6656e+14 | 0.00044431 | 2250.5513 | 2.6656e+14 | 0.54028 | 0.25292
0.01055 | 0.18549 | 0.0018048 | 2.5728e+14 | 0.00042852 | 2333.4894 | 2.5728e+14 | 0.53873 | 0.25311
0.01065 | 0.17917 | 0.0013805 | 2.4846e+14 | 0.00051881 | 1927.3762 | 2.4846e+14 | 0.53721 | 0.2019
0.01075 | 0.17264 | 0.0018107 | 2.3935e+14 | 0.00040147 | 2490.7433 | 2.3935e+14 | 0.53559 | 0.25134
0.01085 | 0.16646 | 0.001628 | 2.3074e+14 | 0.00042598 | 2347.4603 | 2.3074e+14 | 0.53401 | 0.22835
0.01095 | 0.16118 | 0.0013024 | 2.2337e+14 | 0.00050106 | 1995.7048 | 2.2337e+14 | 0.53261 | 0.18794
0.01105 | 0.15531 | 0.0017008 | 2.152e+14 | 0.00038741 | 2581.1833 | 2.152e+14 | 0.53102 | 0.23418
0.01115 | 0.15004 | 0.001338 | 2.0787e+14 | 0.00046341 | 2157.852 | 2.0787e+14 | 0.52954 | 0.1891
0.01125 | 0.14481 | 0.0015646 | 2.0058e+14 | 0.00039579 | 2526.494 | 2.0058e+14 | 0.52803 | 0.21365
0.01135 | 0.14085 | 0.0015611 | 1.9508e+14 | 0.00038691 | 2584.5158 | 1.9508e+14 | 0.52686 | 0.21256
0.01145 | 0.13636 | 0.00115 | 1.8883e+14 | 0.00049014 | 2040.1611 | 1.8883e+14 | 0.52549 | 0.16241
0.01155 | 0.1313 | 0.0016178 | 1.8179e+14 | 0.00035245 | 2837.2359 | 1.8179e+14 | 0.52391 | 0.21745
0.01165 | 0.12769 | 0.0013932 | 1.7677e+14 | 0.00039354 | 2541.0091 | 1.7677e+14 | 0.52274 | 0.18936
0.01175 | 0.12378 | 0.0012608 | 1.7134e+14 | 0.00041883 | 2387.566 | 1.7134e+14 | 0.52145 | 0.17246
0.01185 | 0.11989 | 0.0015377 | 1.6592e+14 | 0.00034245 | 2920.129 | 1.6592e+14 | 0.52013 | 0.20426
0.01195 | 0.11596 | 0.0011818 | 1.6047e+14 | 0.00042041 | 2378.5953 | 1.6047e+14 | 0.51876 | 0.16091
0.01205 | 0.1133 | 0.0014498 | 1.5678e+14 | 0.00034522 | 2896.6488 | 1.5678e+14 | 0.51781 | 0.19146
0.01215 | 0.10965 | 0.0014073 | 1.517e+14 | 0.00034452 | 2902.5858 | 1.517e+14 | 0.51647 | 0.18564
0.01225 | 0.10675 | 0.0011471 | 147678237239640 | 0.00040304 | 2481.0867 | 147678237239640 | 0.51539 | 0.15447
0.01235 | 0.1036 | 0.0015836 | 1.433e+14 | 0 | 0 | 1.433e+14 | 0.51417
0.01245 | 0.10088 | 0.0013619 | 1.3953e+14 | 0 | 0 | 1.3953e+14 | 0.51311

## Report
SEMI PV13 Recombination Lifetime Report
Results:
| Result Description | Symbolic | Value | Units
| Measured effective lifetime | τeff | 363.872 | µs
| Carrier density range | Δn | 1000000000000000 | cm-3
| Measured sheet resistance | Ω/sq | 896.4372 | Ω/sq
| Interpretation parameters: | - | - | -
| Surface recombination | S |  | cm2/s
| Bulk lifetime | τbulk |  | µs
| Emitter saturation current | Joe | 1.1968e-12 | A/cm2
Sample Parameters:
| Parameter Description | Symbolic | Value | Units
| Sample ID | - | nSi_AlOx10B_450 | -
| Sample thickness | w | 0.07 | cm
| Sample temperature | T |  | ⁰C
| Doping concentration | NA (ND) | 1.1092e+14 | cm-3
| Doping type | n/p | n-type
| Surface passivation, front and back | -
| Defect state: | -
| Fe dissociation level | -
| B:O degradation level | -
Measurement type:
Transient
Quasi-steady-state (QSS)
| X | Generalized
Excitation wavelengths/frequencies (file)
| Trapping or DRM correction [if any] | N/A
Instrument parameters:
| Parameter Description |  | Value | Units
| Instrument model & software version |  | WCT-120, Software v4.6.0 | -
| Instrument Serial Number |  | WCT-0318 | -
Light time profile
| Calibration of photogeneration |  | Reference Cell Calibration = 0.008 Optical Constant = 0.7 | V/sun
| Sensor type |  | Eddy Current | -
| Calibration to Δn, including mobility model |  | SEMI PV13 (A,B,C = 0.0002, 0.02769, -0.88568) | -
| Detection area |  | - | -
| Area |  | 12.5664 | cm2
| Number of points |  | 1, center | -
Method of averaging [if any]
4720 Walnut Street, Suite 102, Boulder, CO 80301 USA • +1 303.945.2113 • www.sintoninstruments.com

## Command (hidden)
Parameter | Value | Gain Setting | Flag String
pcgain | 10 | 1 | /pcgain=1 |  | Path name | C:\lifetime\lifetime.exe /subbase=2 /basetime_us=20000
refgain | 0.5 | 20 | /refgain=20 |  | Measure Command | C:\lifetime\lifetime.exe /subbase=2 /basetime_us=20000 /pcgain=1 /refgain=20 /basegain=1 /xfer=125 /flashwait_s=6 /numflashes=1 /pulselength=12500 /cutoff=12500
basegain | 10 | 1 | /basegain=1 |  | Cal Command | C:\lifetime\lifetime.exe /baseonly=true /autoexit=true /basegain=10
Pulse Length | 12500 |  | /pulselength=12500
Sample Time | 100000 |  | /ltransient=100000 |  | Time Scale Flags | /pulselength=12500
# of taus beyond 150mV | 3 |  | /ntransient=3
xfer | 125 |  | /xfer=125
time between flashes | 6 |  | /flashwait_s=6
num flashes for averaging | 1 |  | /numflashes=1
Air Voltage (ZeroCoil) | 1 | 10 | /basegain=10
Cut Off | 12500 |  | /cutoff=12500
Raw Data Graph Points | PC Gain |  | Ref Gain
Time | Upper Bound | Lower Bound | Upper Bound | Lower Bound | Lower Bound Percentage
0 | 10 | 4 | 0.5 | 0.2 | 40%
0.01245 | 10 | 4 | 0.5 | 0.2
Zoom Indicators
Max Voltage | 5.2886 |  | 0.44356
Signal Clipped (Zoom Out) | FALSE |  | FALSE
Signal Low (Zoom In) | FALSE |  | FALSE
Selection Options
TRUE |  | Voltage Range (V) | Gain Setting
FALSE |  | 0.2 | 50
|  | 0.5 | 20
Generalized (1/1) |  | 1 | 10
Generalized (1/64) |  | 2 | 5
QSS |  | 5 | 2
Transient |  | 10 | 1
|  | 20 | -1
n-type |  | 40 | -2
p-type
Sinton
Richter
ON
OFF