
Ming Li, Manus Henry
Department of Engineering Science
University of Oxford
Oxford, OX1 3PJ, UK.
manus.henry@eng.ox.ac.uk
Abstract— Improved signal processing methods for Coriolis Mass Flow Meters provide more accurate measurements and offer a wider range of applications, such as two-phase flow. This paper reviews current signal processing methods for amplitude, frequency and phase difference tracking and analyses their performance in a simulation of two-phase flow conditions.
Keywords—Coriolis Mass Flow Meter, signal processing, dynamic response, time delay, frequency tracking, phase difference tracking, amplitude tracking.
I. Introduction
This research was carried out with the support of Schneider Electric
The Coriolis Mass Flow (CMF) meter, has undergone significant development over the last  20 years [1]. The meter consists of a vibrating flowtube in series with the process piping, and an electronic transmitter that maintains flowtube vibration, performs measurement calculations, and transmits measurement data [2]. The frequency of oscillation (typically 50 Hz - 1 kHz for different flow-tube designs) indicates the density of the process fluid, while the phase difference between the sensor signals indicates the mass flow rate [3]. CMF meters provide high accuracy (e.g. 0.1% of reading for steady flow rates) and good turndown (e.g. 100:1). As digital techniques have developed, the CMF meter has migrated from analogue to digital implementations. Signal processing techniques have evolved to track the amplitude, frequency and phase of the sensor signals for measurement, and to generate a drive signal to maintain flow tube oscillation. Some have suggested the CMF meter is an “almost perfect” flowmeter [4], but limitations have included an inability to maintain flowtube oscillation and/or provide accurate measurements when presented with mixtures of gas and liquid.
Signal processing methods for single phase (i.e. pure gas or liquid) measurement are well established. Techniques in the recent literature [1], [5], [6] include Fourier Transform, Digital Phase Locked Loop, Digital Correlation, Adaptive Notch Filter, and the Hilbert Transform. Two methods not considered are the Sliding Goertzel Algorithm (SGA) [7] and Quadrature Demodulation (QD) [8]. Both have inherent limitations relating to dynamic response that prevent further consideration here. 
The dynamic response [9]–[12] of the CMF meter is a significant issue in two application areas. Conventionally, the response of the meter is important for batching applications such as small volume proving, where industrial standards are based on the response time of mechanical meters [13]. More recently, for CMF meters capable of operating in two-phase or three-phase flow [14], tracking of the rapidly varying amplitude, frequency and phase of the sensor signals is important in order to generate an optimal drive signal [3]. For example, figure 1.1 shows typical sensor signals observed when the flowtube is filled with a single-phase fluid. The amplitude and frequency of the sensor signals are steady. 
[image: image1.png]200
150
100

50

1 Signhal (mV)

-50
-100

Sensor

-150
-200

1

1.0

1 1.02 1.03 1.04 1.05 1.06 1.07 1.08 1.09

Time (s)

200

150

100

50

sor 2 Signal (mV)

-50

5
8 -100

-150

-200

J

1.02 1.03 1.04 1.05 1.06 1.07 1.08 1.09

Time (s)




Fig. 1.1: Sensor signals during steady single phase flow

By contrast, figure 1.2 shows typical sensor signals with a mixture of liquid and gas. Changes in the distribution of liquid and gas in the flowtube lead to variable damping, and rapid changes in the frequency and amplitude of oscillation. The challenge to the amplitude control system is significant, and a fast dynamic response for the measurement of frequency, phase and amplitude should lead to improved flowtube control, as well as measurement, in the presence of two-phase flow.
The purpose of this paper is to evaluate the ability of various signal processing approaches to track rapid changes in the sensor signal properties. With the increasing importance of two-phase flow, this analysis includes the tracking of variable amplitude (often neglected in the prior literature) alongside the calculation of phase difference and frequency.
[image: image2.png]1 Signhal (mV)

Sensor

2 Signal (mV)

Sensor

40

S0 VA AN AN A
ok LU LUV Y

|
_12/\/\/\ PV
FOAL Y VA VALY VAN VYR VAN

-40
15.5 15.51 15652 15563 1554 15565 1556 1557 15568 15.59 15.6
Time (s)

40

o L A A A T A Y B N L
A AN AN A A A

N /\// \\/\ \\// L // \\/\//\ | |
AL A O VA VR VA A

-40
15.5 15.51 15652 15563 1554 15565 1556 1557 15568 15.59 15.6
Time (s)





Fig. 1.2: Sensor signals during two-phase flow
II. Method Analysis
A. Discrete Fourier Transform
A well-established approach for tracking the frequency of a sinusoid is to calculate the signal’s Fourier coefficients and map the results into the frequency domain. The peak in the frequency domain gives the best estimate of the signal’s frequency. Phase difference can be calculated using the Discrete Time Fourier Transform (DTFT) based on the peak frequency. For efficient real-time implementation and/or to reduce spectral leakage, windowing techniques (e.g. rectangular, Hanning) may be applied. Rectangular windowing can be implemented via a recursive calculation to reduce the computational load, while other windowing techniques can reduce spectral leakage.
For rectangular windowing, DFT values at time n can be calculated recursively from the results at time n-1, using


[image: image3.wmf]2/

2/

(,)[(/21)(/21)](1)

(1,)

kjkN

R

jkN

R

DFTnkxnNxnNe

DFTnke

p

p

=+-----

+-

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (1)

Here x(t) is the input signal function, N is the window length and k=0,1,⋯,N/2-1. In [15], Hirokazu et at al. applied the DFT method to track the frequency of a Coriolis flowtube in a commercial implementation.
Once the current frequency has been determined, the phase difference and amplitude can be calculated using the Discrete Time Fourier Transform (DTFT), based on the tracked frequency’s Fourier coefficients. Tu et al. [16] explored the spectral leakage problem and determined the phase of a periodic signal from an arbitrary number of samples without spectral leakage by considering the negative frequency contribution. Shen et al. [17] further improved the recursive DTFT algorithm and increased the accuracy of the phase difference tracking. The equation for their most recent recursive DTFT algorithm is as follows:


[image: image4.wmf]1

,

0

2

(1)

()()

()(1)(2)

(1)

m

mm

m

N

jn

mNm

n

jj

jN

Xxmne

xmxmexme

xmNe

w

ww

w

w

-

-

=

--

--

=+×

=++×++×

+++-×

å

)

))

)

)

L

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (2)

At the next sample time m+1, this becomes:



[image: image5.wmf]1

11

1

1

1

1

1,1

0

1

0

(1)

1

0

(1)

,

()(1)

()()

()

()()

()

[()()()

m

mm

m

mm

m

N

jn

mNm

n

N

jnj

n

jN

N

jnj

n

jN

mNm

j

Xxmne

xmnexme

xmNe

xmnexme

xmNe

XxmxmN

e

w

ww

w

ww

w

w

w

w

+

++

+

+

+

-

-

++

=

-

--

=

--

-

--

=

--

-

=++×

éù

=+×-×

êú

êú

ëû

++×

éù

»+×-×

êú

êú

ëû

++×

=-++×

å

å

å

)

))

)

))

)

)

)

)

11

]

mm

Nj

e

w

++

-

×

)

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (3)

In (2) and (3), 
[image: image6.wmf]m

w

)

 and 
[image: image7.wmf]1

m

w

+

)

 represent the estimated frequencies for the two adjacent sampling points.

To a limited extent, the DFT can overcome the interference of harmonics and noise, is operationally efficient, and is easy to implement in real-time hardware. However, the FFT has inherent performance limitations, such as spectral leakage and the picket fence effect, due to nonsynchronous sampling [18]. Also there is an unavoidable trade-off between time and frequency domain resolution, based on the window length. In two-phase flow situations, the sensor signal varies rapidly, so the measurement needs a fast dynamic response. This requires high time domain resolution which, for the DFT method, will result in decreased frequency domain resolution. Hence the DFT method cannot guarantee high frequency resolution for two-phase flow situations.

For phase difference tracking, the negative frequency-based DTFT algorithm improves the precision but spectral leakage problems still exist with nonsynchronous sampling. Other methods ([19], [20]) have reduced the phase difference error using Rectangular Self-Convolution Windows (RSCWs) but the resulting dynamic response has not been explored. 
B. Digital Phase Locked Loop
Freeman [21] used a Digital Phase Locked Loop (DPLL) to track frequency and phase difference for a CMF meter. As shown in Figure 2.1, the signals are multiplied by a modulation function
[image: image8.wmf]()exp(2)

h

cnjfnt

p

=-


, where  is close to the sensor signal frequency, so that the resulting product is shifted close to zero hertz (DC). Comb filters are used to remove harmonic noise, followed by decimation and down-sampling. Three FIR filters track the down-sampled signals at frequencies –fh /16, 0 and +fh /16 with linear or quadratic interpolation providing a final estimate of the flowtube resonant frequency.

[image: image11.png]HETERODYNE

94

PHASE

ESTIMATE

MIXER COMB
X1(8 @ FILTER
CHANNEL 1 S
[
96

150

/148

100

02
COMB
FILTER

FREQUENCY
ESTIMATE

TRACK QUALITY

TEST
j1 s [134
ADJUST
CPOHI\;I\E%TRE PHASE LOCKED
LOOP
2T TRACKER FREQ. \L
h 142

&>

92

—>




Fig.2.1. Digital Phase Locked Loop diagram [21] 
C. Adaptive Notch Filters
There have been many academic and commercial applications of ANFs to CMF meter frequency tracking, with perhaps the first example described by Derby et al.[22]. Performance improvements are obtained by adopting the Steiglitz-McBride ANF (SMM-ANF) [7], [23]. Typically, a DTFT technique is used to complete the calculation of phase difference and amplitude. Figure 2.2 shows the structure of SMM-ANF algorithm:
[image: image12.png]



Fig. 2.2: Structrue of SMM-ANF [7]
Here the transform function is:


[image: image13.wmf]112

1122

1

()1()

()

()1()()()

m

nk

nkkk

k

Aznzz

Hz

Aznnznz

a

rrar

---

---

=

++

==

++

Õ

)

)

)

)

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (4)

where m is the trap number i.e. the number of  peak frequencies to be tracked. Since the CMF signal has a single dominant frequency, m=1. ρk(n) is the pole contraction factor which determines the bandwidth of the ANF, given by:


[image: image14.wmf]1

2

2()

2cos

1()

k

k

n

BW

n

r

r

-

æö

=

ç÷

ç÷

+

èø

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (5)

The bandwidth is initially broad to capture the frequency of the input signal and is then narrowed to achieve convergence and improved accuracy. Hence ρk(n) is increased over time [24]: 


[image: image15.wmf]00

()(1)(1)

kk

nn

rrrrr

¥

=-+-

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (6)

The weight coefficient 
[image: image16.wmf]()2cos()

kk

nn

aw

=-

))

, with 
[image: image17.wmf]()

k

n

w

)

 the notch frequency estimate of the input frequency, is adjusted by a Recursive Mean Square (RMS) algorithm:


[image: image18.wmf](1)()()()()

kks

nnPnnen

aay

+=-

))

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (7)

where 
[image: image19.wmf]()

n

y

 is the gradient function, 
[image: image20.wmf]()

Pn

 represents the covariance parameters, 
[image: image21.wmf]()

n

l

 is the forgetting factor, and 
[image: image22.wmf]()

s

en

 represents the output of ANF shown in Figure 2.2.: 


[image: image23.wmf]122

()(1)()(1)

()

()

1()()()

sks

k

kkk

enynnen

n

n

nnznz

r

y

a

rar

--

¶---

==

¶

++

)

)

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (8)



[image: image24.wmf]2

(1())(1)

()

()()(1)

nPn

Pn

nnPn

l

ly

--

=

+-

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (9)



[image: image25.wmf]00

()(1)(1)

nn

lllll

¥

=-+-

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (10)

When implementing the SMM-ANF algorithm in the simulation described below, the initial values are:



[image: image26.wmf](

)

(

)

(

)

(

)

(

)

(

)

0

0

2,120.9,0.995,

12100,120.9,0.995

s

fkHz

PP

rrrr

llll

¥

¥

=====

======

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (11)

There are limitations on how well this technique can track signals with rapidly changing amplitude, as occurs in two-phase flow. Recently, Yin et al.[25] developed the Amplitude Adaptive Notch Filter (AANF) to enable the method to work with an amplitude varying signal. However, AANF was designed for Grid Signal Processing, and the application of AANF in CMF signal processing remains unexplored. 

D. Digital Correlation

Wray et al. [26] describe a correlation method for finding the phase difference α between sensor signals, A and B. Given:



[image: image27.wmf]sin

Aat

w

=

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (12)



[image: image28.wmf]sin()

Bbt

wa

=+

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (13)

Then the cross product 
[image: image29.wmf]AB

Ä

  can be written as:



[image: image30.wmf]2

0

sinsin()

ABabttdt

p

wwa

Ä=+

ò

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (14)

Thus the phase difference can be represented as:



[image: image31.wmf]2

AB

ab

a

Ä

=´

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (15)

Further techniques are required to track frequency and amplitude. Although this technique was originally developed for analogue signal processing, Ruoff et al. [5] discuss its merits in the context of digital processing.
E. Hilbert Transform

Duffill et al. [27] applied the Hilbert transform to form an analytic signal for CMF tracking. FIR Hilbert filtering of the sensor signals creates Hilbert pairs. The original left and right pickoff (i.e. sensor) input signals are defined as:


[image: image32.wmf]cos(/2)

orilpo

LPOAwt

q

=+

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (16)



[image: image33.wmf]cos(/2)

orirpo

RPOAwt

q

=-

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (17)

where θ is the phase difference and w is the frequency in radians per second. The Hilbert filtered signals are combined with the original sensor signals to yield:



[image: image34.wmf](/2)

[cos(/2)sin(/2)]

lpo

jwt

lpo

LPOAwtjwt

Ae

q

qq

-

=-+-

=

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (18)

with the equivalent calculation for RPO. Taking the conjugate of the left input Hilbert pairs and multiplying by the right input Hilbert pairs yields the following equation:



[image: image35.wmf](/2)(/2)

jwtjwt

lporpo

j

lporpo

LPORPOAeAe

AAe

qq

q

--+

´=´

=´

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (19)

Assuming the left and right sensor signals have equal amplitude simplifies this to:



[image: image36.wmf]2

(cossin)

LPORPOAj

qq

´=+

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (20)

The phase difference can be obtained using:



[image: image37.wmf]arg()

LPORPO

q

=´

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (21)

The amplitude is calculated from the Hilbert pairs:



[image: image38.wmf]()()

lporpo

AAabsLPOabsRPO

===

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (22)

The frequency is obtained based on the phase change between successive samples: 


[image: image39.wmf]1

1

(/2)(/2)

1

()

2

nn

nn

jwtjwt

n

nlpolpo

jwtwt

lpo

LPOLPOAeAe

Ae

qq

-

-

---

-

-

´=´

=

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (23)

The frequency (in radians) of the left input signal is given by:



[image: image40.wmf]1

1

arg()

n

nnn

wtwtLPOLPO

-

-

-=´

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (24)

Converting into Hz:



[image: image41.wmf]1

()

2

nn

lpo

wtwtFs

f

p

-

-´

=

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (25)

where Fs is the sampling rate of the input signal. Peng et al. [6] claim the Hilbert transform is not suitable for frequency tracking, while Tu et al. [7] use ANF for frequency tracking, and the Hilbert transform for phase difference calculation. However, as shown in the equations above and the results below, the Hilbert technique can be adapted to track frequency, amplitude and phase difference simultaneously. 
III. Simulation results
Previous authors have not discussed the most effective signal processing technique for dealing with mixtures of liquid and gas. The main challenge here is the dynamic response of the signal processing algorithm to deal with rapid and simultaneous changes in amplitude, frequency and phase difference. The importance of fast dynamic response for single phase applications has been thoroughly discussed in [9]–[12].
The authors have developed a set of benchmark simulations to evaluate CMF signal processing techniques against a range of scenarios, entailing linear, sinusoidal and/or random changes in amplitude, frequency and/or phase difference, where such parameter changes may occur singly or in combination. These benchmarks are based on our extensive experience of using CMF meters in two-phase and three-phase applications [14]. The full of benchmark tests will be discussed in future journal papers. Here, only a single test is discussed, which extends an example previously described in the literature [7].
Three of the methods described above have been implemented in simulation, and applied to sensor signals with randomly varying amplitude, frequency and phase difference. The standard deviation of the error for each parameter value has been calculated, and the results compared for each method. Time-varying sensor signals are generated using an amended version of the random walk model from [7], as follows:



[image: image42.wmf]111

222

()()sin[2()/()/2]()

()()sin[2()/()/2]()

()(1)()

()(1)()

()(1)()

se

se

AA

ff

ynAnnfnfnen

ynAnnfnfnen

AnAnen

fnfnen

nnen

ff

pfs

pfs

s

s

ffs

=++×

=-+×

=-+×

=-+×

=-+×

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (26)

Here y1 and y2 are the simulated sensor signals, with common but time varying amplitude A, frequency f, and phase difference (. The sampling update rate for the simulation, fs, is 2 kHz. 
[image: image43.wmf]1

()

en

, 
[image: image44.wmf]2

()

en

, 
[image: image45.wmf]()

A

en


, 
 and 
 are un-correlated white noise processes, which, like using a common amplitude for the two sensor signals, represents a simplification of actual two phase flow conditions. With two-phase flow there is likely to be some correlation between the true mass flow and density values – for example, when a slug of liquid passes through the flowtube both parameter values should increase together. , 
[image: image52.wmf]2

e

s

, 
[image: image53.wmf]A

s

, 
[image: image54.wmf]f

s

 and 
[image: image55.wmf]f

s


random noise. For this simulation, initial values were as follows: A(0) = 0.06V, f(0) = 90Hz, ((0) = 2°, while the noise magnitudes were as follows:  are multiplying factors for the , 
[image: image59.wmf]4

2

10

e

s

-

=

, 
[image: image60.wmf]3

10

A

s

-

=

, 
[image: image61.wmf]4

10

f

s

-

=

 and 
[image: image62.wmf]4

10

f

s

-

=

. Note that a warmup period of approximately 500 samples was used to allow the filters to initialise, and this results in an offset from the initial value at the reported t = 0s in the plots shown below. 
In [7], the simulations reflected single phase conditions. Accordingly, the random variations in frequency and phase difference were small, with maximum variations of 4e-2 Hz and 3e-1 degrees respectively, and no variation in signal amplitude. In the current simulation, as shown in Figure 3.1, the true frequency varies by approximately 10Hz, the phase difference by 6e-1 degrees, and the amplitude by about 4e-2 V, which is approximately ±50% of its mean value.

Results are presented here for the following methods:

· ANF and Hilbert for frequency tracking;
· DTFT (ANF based frequency) and Hilbert for phase difference tracking; and
· DTFT (ANF based frequency) and Hilbert for amplitude tracking.
Note these implementations are our own, but based on the literature cited above. Figure 3.2 - 3.4 show the tracking performance for frequency, phase difference and amplitude respectively. In each case the tracked values are shown in the upper graph, with the residual errors below. The frequency tracking errors are high bandwidth for the Hilbert method and low bandwidth for the ANF method. For phase difference, the Hilbert method shows good tracking with some time delay, while the DTFT  results appear dominated by the frequency errors  of the ANF upon which it relies. Similarly, for amplitude, the DTFT method shows large errors while the Hilbert method tracks relatively well with some time delay. 
[image: image63.png]Amplitude(V)

Frequency(Hz)

Phase Difference(degree)

3 38R

NN
NEEEES

0.09
0.08
0.07
0.06
0.05

True input signal frequency

0 01 02 03 04 05 06 o.
Time(s)
True input signal phase difference
0 01 02 03 04 05 06 o.
Time(s)
True input signal amplitude
0 01 02 03 04 05 06 07

Time(s)




Fig. 3.1: Input signal’s parameters
[image: image64.png]Frequency(Hz)

]

Frequency(Hz)

Frequency Tracking Performance

‘True Frequency.
ANF Tracked Frequency
Hiber Tracked Froauency

ATV

Wi

03 04
Time(s)
Frequency Tracking Errors

05 06 o7

ANF Tracked Frequency Error
Hibert Tracked Freguency Error

Time(s)




Fig. 3.2: Frequency tracking performance

In general terms, therefore, the rapidly changing errors for frequency tracking for the Hilbert method do not result in large amplitude and phase difference errors, while conversely, the slowly changing errors from the ANF method lead to poorer tracking for amplitude and phase when using DTFT. Table I shows the standard deviation for each method and parameter, and the Root Mean Squared Error (RMSE), calculated using:



[image: image65.wmf]^

2

1

1

(()())

n

i

RMSEYnYn

n

=

=-

å

 
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (27)

where 
[image: image66.wmf]()

Yn

 and 
[image: image67.wmf]^

()

Yn

 are the true and estimated values.
[image: image68.png]Phase Difference(degree)

Phase(degree)

Phase Difference Tracking Performance

- [
> » v N o®

,,

True Phase Difference
DTFT(ANF based) Tracked Phase Difference
Hibert Tracked Phase Difference

Time(s)
Phase Difference Tracking Errors

DTFT(ANF based) Tracked Phase Difference Error|
Hibert Tracked Phase Difference Error





Fig. 3.3: Phase difference tracking performance
[image: image69.png]Amplitude Tracking Performance

009 True Ampltude
DTFT(ANF based) Tracked Amplitude i
S 008 ) piber Tracked Ampliude
Foorl ]
2006t 1
£
< 005 0
0.04 1
0 01 02 03 04 05 08 07
Time(s)
Amplitude Tracking Errors
002
s
g o
3002
£
<
04| DTFT(ANF based) Tracked Amplitude Error
Hibert Tracked Amplitude Error
006 N N N . .

0 01 02 03 04 05 06 07
Time(s)




Fig. 3.4: Amplitude tracking performance
The std and RMSE result for frequency are broadly similar for the ANF and Hilbert methods. For phase difference and amplitude, the stds of the Hilbert method are 2-3 times smaller than those for the ANF/DTFT hybrid method. Similarly the RMSE for Hilbert is 3-6 times smaller.
TABLE I.  Quantized Tracking Performance 

	Method and Parameter
	ANF
	DTFT
	Hilbert

	Frequency RMSE (Hz)
	1.0500e+1
	N/A
	1.0136e+1

	Phase Difference RMSE (°)
	N/A
	1.2440e+0
	2.1770e-1

	Amplitude RMSE (V)
	N/A
	2.4073e-2
	7.8850e-3

	Frequency std (Hz)
	1.0493e+1
	N/A
	1.0098e+1

	Phase Difference std (°)
	N/A
	1.2389e+0
	3.1670e-1

	Amplitude std (V)
	N/A
	1.9967e-2
	7.8304e-3


More generally, the advantages and disadvantages the various signal processing techniques discussed in this paper are summarized in Table II as follows:
TABLE II.  Advantages and Disadvantages of Each Method
	Method
	Advantages
	Disadvantages

	Discrete Fourier Transform
	Track all parameters simultaneously (using ANF for frequency tracking); computationally efficient in recursive algorithmic form.
	Spectrum leakage; limited dynamic response; trade-off between time and frequency resolution.

	Digital PLL
	Tracks all parameters simultaneously; suppresses harmonic interference.
	Limited dynamic response; result update once per cycle; limited tracking range.

	Adaptive Notch Filter
	Good ability to suppress noise.
	Limited dynamic response; slow convergence; poor dealing with amplitude change; trade-off between convergence speed and precision.

	Digital Correlation
	Effectively removes random noise; computationally efficient.
	Cannot suppress harmonic noise; result update once per cycle.

	Hilbert Transform
	Track all parameters simultaneously; computationally efficient;

high precision tracking
	Sensitive to random and harmonic noise; needs strict pre-filter; limited dynamic response


IV. Conclusion
With the increasing computation power available at low cost and with low power consumption, there are continuing opportunities to improve the signal processing techniques used for CMF metering. An important future area of CMF meter application is two-phase flow, where all of the measurement parameters are subject to large and rapid variations. Previous research into signal processing techniques has typically excluded consideration of these difficult conditions – specifically, it has commonly been assumed that the amplitude of the sensor signals is constant. This paper has discussed several current methods, and has evaluated their performance in a simulation of two-phase flow conditions, where frequency, amplitude and phase difference have all varied simultaneously. 

Future work will extend the set of benchmark simulations, and develop new signal processing techniques suited to the challenging conditions of two-phase flow measurement.

References
[1]
T. Wang and R. Baker, “Coriolis flowmeters: a review of developments over the past 20 years, and an assessment of the state of the art and likely future directions,” Flow Meas. Instrum., vol. 40, pp. 99–123, Dec. 2014.

[2]
M. Henry, “Self-validating digital Coriolis mass flow meter,” Comput. Control Eng. J., vol. 11, no. 5, pp. 219–227, Oct. 2000.

[3]
M. Zamora and M. P. Henry, “An FPGA Implementation of a Digital Coriolis Mass Flow Metering Drive System,” IEEE Trans. Ind. Electron., vol. 55, no. 7, pp. 2820–2831, Jul. 2008.

[4]
R. J. Reizner, “Coriolis - The almost perfect flow meter,” Comput. Control Eng. J., vol. 14, no. 4, pp. 28–33.

[5]
J. Ruoff, W. Gauchel, and H. Kück, “Advances in Signal Acquisition and Signal Processing of Coriolis Flow Meters,” Procedia Eng., vol. 87, pp. 1585–1588, 2014.

[6]
P. Yi, Y. Tu, M. Xie, and T. Shen, “Analysis of phase difference tracking methods for signal of coriolis mass flowmeter,” in Proceedings of the 10th World Congress on Intelligent Control and Automation, 2012, pp. 4368–4373.

[7]
 Y. Tu,  H. Yang,  H. Zhang, and  X. Liu, “CMF Signal Processing Method Based on Feedback Corrected ANF and Hilbert Transformation,” Meas. Sci. Rev., vol. 14, no. 1, 2014.

[8]
A. Svete, J. Kutin, G. Bobovnik, and I. Bajsić, “Theoretical and experimental investigations of flow pulsation effects in Coriolis mass flowmeters,” J. Sound Vib., vol. 352, pp. 30–45, Sep. 2015.

[9]
C. Clark and R. Cheesewright, “Experimental determination of the dynamic response of Coriolis mass flow meters,” Flow Meas. Instrum., vol. 17, no. 1, pp. 39–47, Mar. 2006.

[10]
C. Clark, R. Cheesewright, S. Wang, M. Henry, M. Zamora, and M. Tombs, “A radically new dynamic response capability for Coriolis flow meters,” Sensors Actuators A Phys., vol. 123–124, pp. 54–62, Sep. 2005.

[11]
C. Clark, R. Cheesewright, and S. Wang, “Prediction of the Dynamic Performance of Fast Response Coriolis Meter Systems,” IEEE Trans. Instrum. Meas., vol. 57, no. 1, pp. 95–99, Jan. 2008.

[12]
M. P. Henry, C. Clark, M. Duta, R. Cheesewright, and M. Tombs, “Response of a Coriolis mass flow meter to step changes in flow rate,” Flow Meas. Instrum., vol. 14, no. 3, pp. 109–118, Jun. 2003.

[13]
M. Tombs, M. Henry, F. Zhou, R. M. Lansangan, and M. Reese, “High precision Coriolis mass flow measurement applied to small volume proving,” Flow Meas. Instrum., vol. 17, no. 6, pp. 371–382, Dec. 2006.

[14]
M. Henry, M. Tombs, M. Zamora, and F. Zhou, “Coriolis mass flow metering for three-phase flow: A case study,” Flow Meas. Instrum., vol. 30, pp. 112–122, Apr. 2013.

[15]
H. Kitami and H. Shimada, “Signal processing method, signal processing apparatus, and coriolis flowmeter,” U.S. Patent 8676517, 18-Mar-2014.

[16]
Y. Tu and H. Zhang, “Method for CMF Signal Processing Based on the Recursive DTFT Algorithm With Negative Frequency Contribution,” IEEE Trans. Instrum. Meas., vol. 57, no. 11, pp. 2647–2654, Nov. 2008.

[17]
T. Shen, Y. Tu, and H. Zhang, “A novel time varying signal processing method for Coriolis mass flowmeter.,” Rev. Sci. Instrum., vol. 85, no. 6, p. 065116, Jun. 2014.

[18]
H. Wen, J. Zhang, Z. Meng, S. Guo, F. Li, and Y. Yang, “Harmonic Estimation Using Symmetrical Interpolation FFT Based on Triangular Self-Convolution Window,” IEEE Trans. Ind. Informatics, vol. 11, no. 1, pp. 16–26, Feb. 2015.

[19]
T. Shen, Y. Tu, M. Li, and H. Zhang, “A new phase difference measurement algorithm for extreme frequency signals based on discrete time Fourier transform with negative frequency contribution.,” Rev. Sci. Instrum., vol. 86, no. 1, p. 015104, Jan. 2015.

[20]
Y. Tu, T. Shen, H. Zhang, and M. Li, “Two New Sliding DTFT Algorithms for Phase Difference Measurement Based on a New Kind of Windows,” Meas. Sci. Rev., vol. 14, no. 6, pp. 350–356, Jan. 2014.

[21]
B. S. Freeman, “Digital phase locked loop signal processing for coriolis mass flow meter,” U.S. Patent 5804741, 08-Sep-1998.

[22]
H. V. Derby, T. Bose, and S. Rajan, “Method and apparatus for adaptive line enhancement in Coriolis mass flow meter measurement,” U.S. Patent 5555190, 10-Sep-1996.

[23]
H. Y. Yang, Y. Q. Tu, and H. T. Zhang, “A Frequency Tracking Method Based on Improved Adaptive Notch Filter for Coriolis Mass Flowmeter,” Appl. Mech. Mater., vol. 128–129, pp. 450–456, Oct. 2011.

[24]
Wei Ni and Kejun Xu, “An improved signal processing method for Coriolis mass flowmeter based on time-varying signal model,” in Fifth World Congress on Intelligent Control and Automation (IEEE Cat. No.04EX788), 2004, vol. 4, pp. 3787–3790.

[25]
G. Yin, L. Guo, and X. Li, “An Amplitude Adaptive Notch Filter for Grid Signal Processing,” IEEE Trans. Power Electron., vol. 28, no. 6, pp. 2638–2641, Jun. 2013.

[26]
T. Wray, “Coriolis flow meter with adjustable excitation phase,” U.S. Patent 6332366, 25-Dec-2001.

[27]
G. R. Duffill, S. M. Jones, and A. T. Patten, “Meter electronics and methods for determining a liquid flow fraction in a gas flow material,” U.S. Patent 7974792, 05-Jul-2011.



_1505128823.unknown

_1505226714.unknown

_1505232213.unknown

_1505234480.unknown

_1505238476.unknown

_1512298830.unknown

_1505236832.unknown

_1505236881.unknown

_1505236873.unknown

_1505235724.unknown

_1505233193.unknown

_1505233590.unknown

_1505233624.unknown

_1505232764.unknown

_1505232115.unknown

_1505232169.unknown

_1505231957.unknown

_1505138436.unknown

_1505139986.unknown

_1505224850.unknown

_1505226633.unknown

_1505225846.unknown

_1505140050.unknown

_1505224783.unknown

_1505140057.unknown

_1505140037.unknown

_1505138774.unknown

_1505139901.unknown

_1505139940.unknown

_1505139967.unknown

_1505139909.unknown

_1505138801.unknown

_1505138756.unknown

_1505138638.unknown

_1505128836.unknown

_1505128847.unknown

_1505128977.unknown

_1505128850.unknown

_1505128843.unknown

_1505128827.unknown

_1505128783.unknown

_1505128809.unknown

_1505128816.unknown

_1505128820.unknown

_1505128812.unknown

_1505128794.unknown

_1505128798.unknown

_1505128786.unknown

_1505128765.unknown

_1505128775.unknown

_1505128779.unknown

_1505128769.unknown

_1505128758.unknown

_1505128762.unknown

_1505128750.unknown

