Estimating radius of bead trajectories from optical trap (632nm)

On the microscope:

- Prepare 1/1000 1um bead slide

- Open calibration program (located in Teuta’s or Lo’s folder depending on iris settings)

- Align bead in trap centre

  - SR = 2000

  - Align with 200 pts to display and 500 pts to av

  - Press stop program button and immediately change the parameters:

    - Pts to Av (speed of nanoncube) = 500 (and also record one file 200 + 1000)

    - Pts to display = 1

    - File length = 0.5sec

  - Wait while the program moves the stage in a grid so that the signal from the bead is mapped in 1D (so known µm moved by the stage can be equated to signal (derived from 4 voltages of QPD)) ie output is X (µm) vs x (relative).  Factors effecting relationship = sixe of laser focus (iris), wavelength, bead size).

  - The stage will move from 19.1 to 20.9 – make sure you record for long enough to repeat one cycle (with minimal drift) eg 0.5sec recording is ok (takes ~5-10mins in real time)

- Output the QuView file to text files (use pts = 1000 works ok)

- Now change slides and record rotating beads as normal

Analysis1:

- Open the file from the first grid run

- Create an ‘x’ file (see attached) by centring the range (19.1-20.9) around zero (also must invert numbers to give correct output (multiply by -1))

- Open “fit scan data--Richards-program-reading-in-Teutas-files-not-rearanging-order”

- Count # of pts in the grid eg 19 and enter in n#x, n#y, n#z (z doesn’t matter = 1; check that y number of pts = x (doesn’t ave to be the same but must be correct)

- Work out the number of pts of x and y for one cycle (eg 19x19 = 361) and copy and paste the x-y columns containing this data into a new text file

- Do for the remaining runs (eg twice more maybe)

- Run the program (“fit scan data--Richards-program-reading-in-Teutas-files-not-rearanging-order”) with an initial linear value of 0.22 or 0.21.

- Set miny, minx, nx, ny so that you are only fitting in the linear range (ie in-between the two populated regions)

- Play around until you get a nice linear fit

- Then go into the block diagram and double-click on the vi highlighted here:

- Press ok once and then again.

- Go back to the original program (“fit scan data--Richards-program-reading-in-Teutas-files-not-rearanging-order”) and make a note (or copy/paste) the ‘x poly fit parameters’ and the ‘y poly fit parameters’,  4 values each

- Repeat the process for other grid runs and compare values, choose one/av and enter into program “calibrated speed by r,theta-4Voltages-Rh-with-radius”

Analysis2

- With the parameters set in “calibrated speed by r,theta-4Voltages-Rh-with-radius”, go through the data set of spinning beads as normal, click polynomial to get ‘corrected’ x and y values (in microns not rel units), use 200median filter and 100bins

- Record average radius and it’s standard deviation for each file of each cell in an excel document, then make averages

- For x-ys plots use “calibrated speed by r,theta-4Voltages-Rh-with-radius FOR XY FIGURE” (just has nice x-y scale etc for easy export).