## Slide 1: PLQE Sphere Calibration
Manuel Kober-Czerny
04-May-2021

## Slide 2: Some Information
Why the reference lamp is important
Use the reference lamp (OceanInsight HL-3P-INT-CAL). The type of lamp is important, since it has a built-in diffuser and doesn’t require an optical fiber to get the light into the sphere. It is therefore much better for uniform illumination of this type of sphere.
Why do we need to calibrate
In short: The spectrometer can only measure ‘counts’, but the PLQE is defined as ratio of photons, so we need to translate this. Therefore we use a calibrated lamp with a known energy to transform counts -> energy. Then we can use the relation of wavelength and energy to get the number of photons from that.
How to calibrate the sphere
As for any normal measurement the ‘non-linearity correction’ should be turned on!
After having warmed up for 15 min. the reference lamp is placed at different entrance ports and the spectrum is measured. Make sure that the detector is not saturated and also measure a background for each of the integration times. Use Bernard’s script to calculate the correction files from this measurement.

## Slide 3: Reference Lamp Measurements
Calibration of the Sphere (Red Fiber)
To check linearity, the reference spectrum of the lamp was measured for different integration times.
Conclusion: The detector response is perfectly linear.

## Slide 4: Reference Lamp Measurements
Calibration of the Sphere (50 um Steel Fiber)
To check linearity, the reference spectrum of the lamp was measured for different integration times.
Conclusion: The detector response is perfectly linear.

## Slide 5: Reference Files of Sphere
Calibration of the Sphere (Red Fiber)
Bernard W. calculated the correction files using the script in the PLQE folder. Even on a logarithmic scale the correction files for different integration times overlapped perfectly, so they were averaged into one file. One of these files was generated for each exit port of the sphere to check uniformity. The results are shown. The Correction value for the top port is lower, most probably because the hole is larger and more light is lost. The correction values obtained from the PLQE and ELQE side overlap
nicely. Since they are mainly used, they are averaged again to yield a single correction file.
This file is cal_QEpro_Red_200.txt.

## Slide 6: Reference Files of Sphere
Calibration of the Sphere (50 um Steel Fiber)
Bernard W. calculated the correction files using the script in the PLQE folder. Even on a logarithmic scale the correction files for different integration times overlapped perfectly, so they were averaged into one file. One of these files was generated for each exit port of the sphere to check uniformity. The results are shown. The Correction value for the top port is lower again, most probably because the hole is larger and more light is lost. The correction values obtained from the PLQE side are strange and might have to be remeasured one day. The EQE side is as expected and hence used for the correction file.
This file is cal_QEpro_Steel-50um_200.txt.

## Slide 7: Comparing Estimated and Measured Power
Some Initial Notes on the Power Estimation
Since the sphere is now calibrated properly and the signal from the spectrometer is translated into W/(nm s) one can easily (in theory) integrate the empty.txt measurement and multiply by the integration time to get an estimate for the laser power in W.
First, I wanted to check, how good this estimate is, so I compared the power estimated from the empty measurement to the reading on a power meter placed in front of the sphere. (Please note that the power meter was also placed behind the sphere to check, if there is a loss of power, but the values are the same. So we can assume that all the light enters the sphere)
Still, the power was underestimated by about 20%. This seemed to be independent of the laser source used. A few things that I noticed, when trying to find out why:
The estimated power dropped, when the front cover was removed, as a lot of light could then leave the sphere. This cover has to be removed though to use the calibration lamp, so there is an error here.
The estimated power went up, when the top cover was removed, which was strange to me, but might be a result of its dirtiness and the high absorption of the metallic parts
I moved the laser spot on the back of the sphere (with the front cover removed) by up to a cm up, down, right and left to see no change in the estimated power. So the scattering surface does not seem to be the major problem.
Using a diffuser in place of the sample would increase the estimated power, so the high power density of the laser does seem to be a problem.

## Slide 8: Comparing Estimated and Measured Power
Estimating the Error of the Power Estimation
To get a grasp of what was happening, I measured the power outside the sphere (power meter reading) and then used three methods to acquire empty measurements:
Using a clean cover instead of the sample holder.
With the sample holder in place (as usually done)
With a diffuser in place of the sample in the sample holder.
All measurements were done for the red and 50 um steel fiber.
A few things to note:
The estimations for the power for both fibers align perfectly, meaning that they are not the sources of error
For the 532 nm laser the diffuser in place of the sample yields an almost 1:1 agreement with the power meter reading
In conclusion: The high power density seems to be the problem that could be addressed by changing the sphere geometry (additional baffles in front of the detector) or by a correction factor. The latter is easier implementable for now.

## Slide 9: Comparing Estimated and Measured Power
Estimating the Error of the Power Estimation
To get a grasp of what was happening, I measured the power outside the sphere (power meter reading) and then used three methods to acquire empty measurements:
Using a clean cover instead of the sample holder.
With the sample holder in place (as usually done)
With a diffuser in place of the sample in the sample holder.
All measurements were done for the red and 50 um steel fiber.
A few things to note:
The estimations for the power for both fibers align nicely, meaning that they are not the sources of error
None of the measurements yield a 1:1 agreement between the estimation and the reading. A reason could be the high correction value at this wavelength, which also increases the error.
In conclusion: A correction factor may be the best option here.

## Slide 10: Power Correction Factors
In Conclusion
The best results would be obtained by measuring the empty sphere with a diffuser in place, but it is not the most practical, especially in terms of automation. The lowest-power measurements were excluded, as they are at the edge of the power meter noise anyways. So, two correction factors are implemented:
Laser | Correction Factor
405 nm | 1.31 +/- 0.09 (Error = 7%)
532 nm | 1.28 +/- 0.05 (Error = 4%)
As a correction factor for the sphere was expected (independent of wavelength), the average was taken to be:
1.30 +/- 0.05 (4% Error)

## Slide 11: Final Power Estimation
Final Comments
The correct power can now be estimated from the empty measurement, when the sample holder is in place. The error of the slope is accurate to the third digit.
Looking at the spread of the data, a generous 10% error in power estimation can be assumed for publications.
This shouldn’t have a major impact on any power-dependent PLQE measurements or recombination constant fittings.